## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 19346-4
Model Number |  | 1281 |  |  | Test date |  | 24-11-25 11:16:02
Meter Info |  |  |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1,1,0,0,0 |  |  | CERT_AMB?
Firmware |  | 890144-03.12 |  |  | TEMP?
Date of Baseline Cal |  |  |  |  | TEMP? BASE
Date of Certified Cal |  |  |  |  | TEMP? CERTIFIED
CNFTST? |  | 0 |  |  | STANDBY?
DUT Confidence level
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 191024 |  |  | MFC last calibrated |  | 37 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 191024 |  |  | MFC since DCV Zero |  | 37 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529912686322891E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306135899178203E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039668828305E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000065537234996E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996275049758838E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6073 days 5.02 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 2 days 1.62 hours |  |  | CAL CONST,DACLIN |  | 0
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.361 | 7.5e-07 |  |  | N/A |  |  | 0.089088
0.0VDC | 0 | 0.07 | 7.5e-07 |  |  | N/A |  |  | 0.052308
0.0VDC | 0 | -1.6 | 7.5e-07 |  |  | N/A |  |  | 0.27839
0.0VDC | 0 | 0 | 7.5e-07 |  |  | N/A |  |  | 2.2423
0.0VDC | 0 | -90 | 7.5e-07 |  |  | N/A |  |  | 16.6667
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.010001 | 81.5 |  |  | 61 |  |  | 2.4724
0.02 | 100mV | 0.020001 | 44 |  |  | 26.6 |  |  | 0.55169
0.05 | 100mV | 0.05 | 21.5 |  |  | 5.2 |  |  | 0.25815
0.1 | 100mV | 0.1 | 14 | 0.099999 | 0.1 | -1.52 | 3 | -10.6161 | 0.13274
-0.01 | 100mV | -0.0099992 | -81.5 |  |  | -78.5 |  |  | -2.7684
-0.02 | 100mV | -0.019999 | -44 |  |  | -42.6 |  |  | -0.37751
-0.05 | 100mV | -0.049999 | 21.5 |  |  | -22.16 |  |  | -0.23367
-0.1 | 100mV | -0.099998 | 14 | -0.099999 | -0.1 | -15.04 | 3 | -105.0439 | -0.11219
0.1 | 1V | 0.1 | 14 |  |  | -5 |  |  | 0.31136
0.2 | 1V | 0.2 | 10.25 |  |  | -5.7 |  |  | 0.1236
0.5 | 1V | 0.5 | 5.9 |  |  | -2.2 |  |  | 0.10745
1 | 1V | 1 | 4.7 | 1 | 1 | -2.26 | 1.5 | -45.8087 | 0.085
-0.1 | 1V | -0.099999 | 14 |  |  | -11.7 |  |  | -0.44191
-0.2 | 1V | -0.2 | 10.25 |  |  | -9.8 |  |  | -0.12748
-0.5 | 1V | -0.5 | 5.9 |  |  | -4.04 |  |  | -0.066416
-1 | 1V | -1 | 4.7 | -1 | -1 | -3.74 | 1.5 | -75.8073 | -0.062071
1 | 10V | 1 | 4.7 |  |  | -4.5 |  |  | 0.25
2 | 10V | 2 | 4.1 |  |  | -3.75 |  |  | 0.091667
5 | 10V | 5 | 4.1 |  |  | -3.02 |  |  | 0.062805
10 | 10V | 10 | 3.8 | 10 | 10 | -2.86 | 1.5 | -70.0064 | 0.028284
19 | 10V | 18.9999 | 3.8158 | 18.9999 | 19.0001 | -2.6526 | 1.5 | -64.6977 | 0.021397
-1 | 10V | -1 | 4.7 |  |  | -0.8 |  |  | -0.40035
-2 | 10V | -2 | 4.1 |  |  | -0.95 |  |  | -0.1523
-5 | 10V | -5 | 4.1 |  |  | -1.12 |  |  | -0.072111
-10 | 10V | -10 | 3.8 | -10 | -10 | -1.31 | 0.5 | -34.1791 | -0.02048
-19 | 10V | -19 | 3.8158 | -18.9999 | -19.0001 | -1.2474 | 1.5 | -30.4233 | -0.013674
10 | 100V | 10 | 3.8 |  |  | -0.8 |  |  | 0.30822
20 | 100V | 20 | 3.8 |  |  | -1.7 |  |  | 0.091667
50 | 100V | 49.9999 | 7 |  |  | -1.28 |  |  | 0.08413
100 | 100V | 99.9998 | 6 | 99.9994 | 100.0006 | -1.53 | 2 | -24.1914 | 0.031002
-10 | 100V | -10 | 3.8 |  |  | -2 |  |  | -0.26977
-20 | 100V | -20 | 3.8 |  |  | -2.05 |  |  | -0.19843
-50 | 100V | -50 | 7 |  |  | -0.88 |  |  | -0.13017
-100 | 100V | -99.9999 | 6 | -99.9994 | -100.0006 | -1.32 | 2 | -20.871 | -0.055252
-100 | 1000V | -100 | 6 |  |  | -0.4 |  |  | -0.26926
-200 | 1000V | -199.9997 | 5.5 |  |  | -1.6 |  |  | -0.13463
-500 | 1000V | -499.9989 | 8.2 |  |  | -2.28 |  |  | -0.049103
-1000 | 1000V | -999.9962 | 7.6 | -999.9921 | -1000.0079 | -3.78 | 2 | -48.0992 | -0.039193
100 | 1000V | 99.9998 | 6 |  |  | -1.5 |  |  | 0.29155
200 | 1000V | 199.9996 | 5.5 |  |  | -1.9 |  |  | 0.1024
500 | 1000V | 499.9988 | 8.2 |  |  | -2.34 |  |  | 0.049441
1000 | 1000V | 999.9961 | 7.6 | 999.9921 | 1000.0079 | -3.91 | 2 | -49.7534 | 0.030414
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99963 | 0.99952 | 85 | 0.99954 | 0.99954 | -105.5393 | 20 | -120.8633 | 4.0776
1.9 Ω | 1.9 | 1.8999 | 85 | 1.8998 | 1.9002 | -65.1587 | 15 | -75.4908 | 1.8937
10 Ω | 10.0003 | 10.0002 | 26 | 10 | 10.0006 | -8.9997 | 5 | -33.9915 | 0.32015
19 Ω | 19.0038 | 19.0037 | 24 | 19.0033 | 19.0044 | -5.42 | 5 | -22.1085 | 0.15291
100 Ω | 100.0049 | 100.0043 | 15 | 100.0031 | 100.0067 | -5.8297 | 3 | -38.11 | 0.027587
190 Ω | 190.0023 | 190.0017 | 15 | 189.9989 | 190.0057 | -3.1158 | 3 | -20.3683 | 0.023307
1 kΩ | 999.9418 | 999.9406 | 11 | 999.9278 | 999.9558 | -1.2201 | 3 | -10.7007 | 0.10249
1.9 kΩ | 1899.9019 | 1899.9003 | 11 | 1899.8753 | 1899.9285 | -0.83689 | 3 | -7.34 | 0.055951
10 kΩ | 10000.397 | 10000.3845 | 9 | 10000.277 | 10000.517 | -1.25 | 3 | -13.1756 | 0.077779
19 kΩ | 18999.633 | 18999.5904 | 9 | 18999.405 | 18999.861 | -2.2421 | 3 | -23.6343 | 0.026316
100 kΩ | 99998.22 | 99998.356 | 11 | 99996.62 | 99999.82 | 1.36 | 5 | 11.2556 | 0.023334
190 kΩ | 189993.74 | 189994.093 | 11 | 189990.7001 | 189996.7799 | 1.858 | 5 | 15.3766 | 0.014679
1 MΩ | 999904.9 | 999894.16 | 16 | 999880.9023 | 999928.8977 | -10.741 | 8 | -60.0441 | 0.096576
1.9 MΩ | 1899848.4 | 1899826.14 | 17 | 1899800.904 | 1899895.896 | -11.7167 | 8 | -62.3618 | 0.028481
10MΩ | 9998508 | 9998137.7 | 33 | 9998058.067 | 9998957.933 | -37.0355 | 12 | -105.472 | 0.065098
19 MΩ | 18999402 | 18998951.6 | 43 | 18998357.03 | 19000446.97 | -23.706 | 12 | -53.1013 | 0.44789
100MΩ(2W) | 100009510 | 99934400 | 110 | 99980507.24 | 100038512.8 | -751.0286 | 180 | -356.0217 | 8.5756
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000135 | 5e-05 |  |  | N/A |  |  | 1.4748e-05
0 Ω | 100 Ω | -9.5e-05 | 5e-05 |  |  | N/A |  |  | 5.4874e-06
0 Ω | 1 KΩ | -8.2e-05 | 5e-05 |  |  | N/A |  |  | 3.1623e-06
0 Ω | 10 KΩ | 0.00118 | 5e-05 |  |  | N/A |  |  | 5.8547e-05
0 Ω | 100 KΩ | 0.0129 | 5e-05 |  |  | N/A |  |  | 0.00031136
0 Ω | 1 MΩ | -0.235 | 5e-05 |  |  | N/A |  |  | 0.0023688
0 Ω | 10 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.026352
0 Ω | 100 MΩ(2W) | -23.5 | 5e-05 |  |  | N/A |  |  | 0.22913
|  | 0 | 5e-05 | [5 empty] | 0
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 1.5662 | 5e-05 |  |  | N/A |  |  | 0.00012468
0 Ω | 100 Ω | 1.5659 | 5e-05 |  |  | N/A |  |  | 6.1594e-05
0 Ω | 1 KΩ | 1.5657 | 5e-05 |  |  | N/A |  |  | 4.4054e-05
0 Ω | 10 KΩ | 1.5666 | 5e-05 |  |  | N/A |  |  | 6.8191e-05
0 Ω | 100 KΩ | 1.5769 | 5e-05 |  |  | N/A |  |  | 0.00038079
0 Ω | 1 MΩ | 1.327 | 5e-05 |  |  | N/A |  |  | 0.0025386
0 Ω | 10 MΩ | 1.55 | 5e-05 |  |  | N/A |  |  | 0.040757
0 Ω | 100 MΩ | -22.1 | 5e-05 |  |  | N/A |  |  | 0.17638
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99988 | 77 | 0.99992 | 1.0001 | -120 | 10 | -154.5463 | 10.0012
2VAC/1kHz | 10 VAC | 1.9999 | 73.5 | 1.9999 | 2.011 | -25 | 10 | -33.7031 | 3.9088
5VAC/1kHz | 10 VAC | 4.9997 | 84 | 4.9996 | 5.0004 | -50 | 10 | -59.1064 | 2.8481
10VAC/1kHz | 10 VAC | 9.9995 | 77 | 9.9992 | 1.006 | -53 | 10 | -68.2579 | 1.4143
12VAC/1kHz | 10 VAC | 11.9994 | 75.833 | 11.9991 | 1.0001 | -52.5 | 10 | -68.6369 | 0.55558
15VAC/1kHz | 10 VAC | 14.9992 | 74.667 | 14.9989 | 1.0057 | -54.6667 | 10 | -72.5661 | 1.0244
19VAC/1kHz | 10 VAC | 18.999 | 73.684 | 18.9986 | 19.0014 | -54.2105 | 10 | -72.9033 | 0.79438
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009959 | 5500 | 0.0009945 | 0.0010055 | -4100 | 20 | -74.545 | 304.4718
0.001 V, 20 Hz |  | 0.0009975 | 5200 | 0.0009948 | 0.0010052 | -2500 | 20 | -48.0766 | 194.8522
0.001 V, 30 Hz |  | 0.0009979 | 5200 | 0.0009948 | 0.0010052 | -2100 | 20 | -40.3843 | 332.3604
0.001 V, 40 Hz |  | 0.0009979 | 5100 | 0.0009949 | 0.0010051 | -2100 | 20 | -41.1762 | 580.9709
0.001 V, 55 Hz |  | 0.0009977 | 5100 | 0.0009949 | 0.0010051 | -2300 | 20 | -45.0977 | 493.0124
0.001 V, 300 Hz |  | 0.000998 | 5100 | 0.0009949 | 0.0010051 | -2000 | 20 | -39.2154 | 301.0647
0.001 V, 1 KHz |  | 0.0009975 | 5100 | 0.0009949 | 0.0010051 | -2500 | 20 | -49.0192 | 601.7358
0.001 V, 10 KHz |  | 0.0009986 | 5100 | 0.0009949 | 0.0010051 | -1400 | 20 | -27.4508 | 668.0184
0.001 V, 20 KHz |  | 0.0009991 | 5100 | 0.0009949 | 0.0010051 | -900 | 20 | -17.6469 | 287.4869
0.001 V, 30 KHz |  | 0.0009992 | 5340 | 0.00099466 | 0.0010053 | -800 | 20 | -14.9812 | 528.2589
0.001 V, 50 KHz |  | 0.0009975 | 5340 | 0.00099466 | 0.0010053 | -2500 | 20 | -46.8162 | 852.1303
0.001 V, 100 KHz |  | 0.0009788 | 8800 | 0.0009912 | 0.0010088 | -21200 | 50 | -240.9052 | 250.2544
0.001 V, 200 KHz |  | 0.0009158 | 16100 | 0.0009839 | 0.0010161 | -84200 | 100 | -522.9713 | 232.3496
0.001 V, 300 KHz |  | 0.0008812 | 16100 | 0.0009839 | 0.0010161 | -118800 | 100 | -737.874 | 260.706
0.001 V, 500 KHz |  | 0.0008323 | 31500 | 0.0009685 | 0.0010315 | -167700 | 200 | -532.3702 | 413.7927
0.001 V, 1 MHz |  | 0.000644 | 33000 | 0.000967 | 0.001033 | -356000 | 300 | -1078.7433 | 268.952
0.01 V, 10 Hz |  | 0.0099726 | 1100 | 0.009989 | 0.010011 | -2740 | 20 | -249.0497 | 24.045
0.01 V, 20 Hz |  | 0.0099923 | 800 | 0.009992 | 0.010008 | -770 | 20 | -96.2199 | 30.6195
0.01 V, 30 Hz |  | 0.0099963 | 800 | 0.009992 | 0.010008 | -370 | 20 | -46.2356 | 13.3383
0.01 V, 40 Hz |  | 0.0099975 | 700 | 0.009993 | 0.010007 | -250 | 20 | -35.6997 | 35.8714
0.01 V, 55 Hz |  | 0.0099981 | 700 | 0.009993 | 0.010007 | -190 | 20 | -27.1318 | 54.5794
0.01 V, 300 Hz |  | 0.0099992 | 700 | 0.009993 | 0.010007 | -80 | 20 | -11.4239 | 71.8466
0.01 V, 1 KHz |  | 0.0099992 | 700 | 0.009993 | 0.010007 | -80 | 20 | -11.4239 | 44.4758
0.01 V, 10 KHz |  | 0.0099993 | 700 | 0.009993 | 0.010007 | -70 | 20 | -9.9959 | 35.3578
0.01 V, 20 KHz |  | 0.0099993 | 700 | 0.009993 | 0.010007 | -70 | 20 | -9.9959 | 37.5674
0.01 V, 30 KHz |  | 0.0099998 | 940 | 0.0099906 | 0.010009 | -20 | 20 | -2.1272 | 58.9503
0.01 V, 50 KHz |  | 0.010003 | 940 | 0.0099906 | 0.010009 | 300 | 20 | 31.9077 | 37.0699
0.01 V, 100 KHz |  | 0.010013 | 1600 | 0.009984 | 0.010016 | 1310 | 50 | 81.8351 | 13.2114
0.01 V, 200 KHz |  | 0.010009 | 2600 | 0.009974 | 0.010026 | 880 | 100 | 33.8211 | 17.6228
0.01 V, 300 KHz |  | 0.010001 | 2600 | 0.009974 | 0.010026 | 90 | 100 | 3.459 | 17.6368
0.01 V, 500 KHz |  | 0.0099071 | 4500 | 0.009955 | 0.010045 | -9290 | 200 | -206.2409 | 19.1811
0.01 V, 1 MHz |  | 0.0086352 | 6000 | 0.0099399 | 0.01006 | -136480 | 300 | -2271.8287 | 33.3184
0.1 V, 10 Hz |  | 0.099728 | 660 | 0.099934 | 0.10007 | -2721 | 20 | -412.0836 | 24.2516
0.1 V, 20 Hz |  | 0.099926 | 300 | 0.09997 | 0.10003 | -745 | 20 | -247.7833 | 3.2039
0.1 V, 30 Hz |  | 0.099964 | 300 | 0.09997 | 0.10003 | -362 | 20 | -120.3994 | 3.2415
0.1 V, 40 Hz |  | 0.099977 | 195 | 0.09998 | 0.10002 | -227 | 20 | -115.8028 | 6.9618
0.1 V, 55 Hz |  | 0.099986 | 195 | 0.09998 | 0.10002 | -145 | 20 | -73.9709 | 4.6644
0.1 V, 300 Hz |  | 0.099995 | 195 | 0.09998 | 0.10002 | -52 | 20 | -26.5275 | 4.3623
0.1 V, 1 KHz |  | 0.099995 | 195 | 0.09998 | 0.10002 | -52 | 20 | -26.5275 | 7.2595
0.1 V, 10 KHz |  | 0.099984 | 195 | 0.09998 | 0.10002 | -165 | 20 | -84.1738 | 5.2687
0.1 V, 20 KHz |  | 0.099969 | 195 | 0.09998 | 0.10002 | -310 | 20 | -158.1447 | 5.3428
0.1 V, 30 KHz |  | 0.099959 | 400 | 0.09996 | 0.10004 | -409 | 20 | -102.1224 | 4.2378
0.1 V, 50 KHz |  | 0.099953 | 400 | 0.09996 | 0.10004 | -469 | 20 | -117.1037 | 4.7045
0.1 V, 100 KHz |  | 0.099967 | 1050 | 0.099895 | 0.10011 | -333 | 50 | -31.6784 | 3.8092
0.1 V, 200 KHz |  | 0.099886 | 1240 | 0.099876 | 0.10012 | -1145 | 100 | -92.0399 | 2.0299
0.1 V, 300 KHz |  | 0.099633 | 1240 | 0.099876 | 0.10012 | -3668 | 100 | -294.8492 | 3.1739
0.1 V, 500 KHz |  | 0.098357 | 1900 | 0.099809 | 0.10019 | -16433 | 200 | -860.1425 | 2.1568
0.1 V, 1 MHz |  | 0.086843 | 4000 | 0.099599 | 0.1004 | -131571 | 300 | -3280.0628 | 2.7679
1 V, 10 Hz |  | 0.99732 | 600 | 0.9994 | 1.0006 | -2682 | 10 | -446.9379 | 11.7576
1 V, 20 Hz |  | 0.99928 | 180 | 0.99982 | 1.0002 | -716 | 10 | -397.1653 | 2.3996
1 V, 30 Hz |  | 0.99966 | 180 | 0.99982 | 1.0002 | -344 | 10 | -190.8169 | 2.2921
1 V, 40 Hz |  | 0.99979 | 77 | 0.99992 | 1.0001 | -215 | 10 | -276.8954 | 2.5501
1 V, 55 Hz |  | 0.99987 | 77 | 0.99992 | 1.0001 | -135 | 10 | -173.8646 | 1.9652
1 V, 300 Hz |  | 0.99996 | 77 | 0.99992 | 1.0001 | -44 | 10 | -56.667 | 2.2792
1 V, 1 KHz |  | 0.99996 | 77 | 0.99992 | 1.0001 | -40 | 10 | -51.5154 | 1.4143
1 V, 10 KHz |  | 0.99985 | 77 | 0.99992 | 1.0001 | -146 | 10 | -188.0313 | 0.86615
1 V, 20 KHz |  | 0.9997 | 77 | 0.99992 | 1.0001 | -301 | 10 | -387.6536 | 1.8339
1 V, 30 KHz |  | 0.99961 | 140 | 0.99986 | 1.0001 | -388 | 10 | -276.4386 | 1.3339
1 V, 50 KHz |  | 0.99956 | 140 | 0.99986 | 1.0001 | -439 | 20 | -310.4199 | 2.0625
1 V, 100 KHz |  | 0.99952 | 310 | 0.99969 | 1.0003 | -484 | 30 | -155.403 | 3.5529
1 V, 200 KHz |  | 0.99941 | 550 | 0.99945 | 1.0006 | -592 | 70 | -106.775 | 1.3953
1 V, 300 KHz |  | 0.99965 | 550 | 0.99945 | 1.0006 | -355 | 70 | -64.029 | 1.2252
1 V, 500 KHz |  | 1.0009 | 1400 | 0.9986 | 1.0014 | 907 | 100 | 64.6211 | 1.48
1 V, 1 MHz |  | 1.0038 | 3000 | 0.99699 | 1.003 | 3819 | 200 | 127.0181 | 4.6578
10 V, 10 Hz |  | 9.973 | 600 | 9.994 | 10.006 | -2697 | 10 | -449.4376 | 22.7335
10 V, 20 Hz |  | 9.9927 | 180 | 9.9982 | 10.0018 | -729 | 10 | -404.3764 | 1.9017
10 V, 30 Hz |  | 9.9964 | 180 | 9.9982 | 10.0018 | -359 | 10 | -199.1374 | 2.5009
10 V, 40 Hz |  | 9.9977 | 77 | 9.9992 | 10.0008 | -229 | 10 | -294.9259 | 2.4043
10 V, 55 Hz |  | 9.9985 | 77 | 9.9992 | 10.0008 | -149 | 10 | -191.895 | 1.6586
10 V, 300 Hz |  | 9.9994 | 77 | 9.9992 | 10.0008 | -57 | 10 | -73.4095 | 1.9651
10 V, 1 KHz |  | 9.9995 | 77 | 9.9992 | 10.0008 | -52 | 10 | -66.9701 | 1.3643
10 V, 10 KHz |  | 9.9992 | 77 | 9.9992 | 10.0008 | -84 | 10 | -108.1824 | 0.88199
10 V, 20 KHz |  | 9.9981 | 77 | 9.9992 | 10.0008 | -194 | 10 | -249.8498 | 1.5814
10 V, 30 KHz |  | 9.997 | 140 | 9.9986 | 10.0014 | -302 | 10 | -215.1661 | 0.7267
10 V, 50 KHz |  | 9.9955 | 140 | 9.9986 | 10.0014 | -451 | 20 | -318.9052 | 0.72681
10 V, 100 KHz |  | 9.9933 | 270 | 9.9973 | 10.0027 | -674 | 30 | -248.1028 | 1.1675
10 V, 200 KHz |  | 9.989 | 670 | 9.9933 | 10.0067 | -1098 | 70 | -162.9934 | 1.4158
10 V, 300 KHz |  | 9.9864 | 670 | 9.9933 | 10.0067 | -1358 | 70 | -201.5893 | 1.9463
10 V, 500 KHz |  | 9.9968 | 1700 | 9.983 | 10.017 | -321 | 100 | -18.8498 | 1.8339
10 V, 1 MHz |  | 10.1272 | 3500 | 9.9649 | 10.0351 | 12723 | 200 | 362.9222 | 6.3227
19 V, 10 Hz |  | 18.9532 | 552.63 | 18.9895 | 19.0105 | -2465.2632 | 10 | -446.0235 | 17.1572
19 V, 20 Hz |  | 18.9872 | 165.79 | 18.9968 | 19.0032 | -675.2632 | 10 | -406.5614 | 1.3795
19 V, 30 Hz |  | 18.9937 | 165.79 | 18.9968 | 19.0032 | -333.1579 | 10 | -200.5872 | 2.6251
19 V, 40 Hz |  | 18.9959 | 73.684 | 18.9986 | 19.0014 | -216.8421 | 10 | -291.6133 | 1.4523
19 V, 55 Hz |  | 18.9973 | 73.684 | 18.9986 | 19.0014 | -143.1579 | 10 | -192.5214 | 1.3363
19 V, 300 Hz |  | 18.9989 | 73.684 | 18.9986 | 19.0014 | -58.4211 | 10 | -78.5657 | 1.9299
19 V, 1 KHz |  | 18.999 | 73.684 | 18.9986 | 19.0014 | -53.1579 | 10 | -71.4877 | 0.87284
19 V, 10 KHz |  | 18.9983 | 73.684 | 18.9986 | 19.0014 | -91.0526 | 10 | -122.4493 | 1.1097
19 V, 20 KHz |  | 18.9961 | 73.684 | 18.9986 | 19.0014 | -206.8421 | 10 | -278.1651 | 3.4419
19 V, 30 KHz |  | 18.9939 | 130.53 | 18.9975 | 19.0025 | -322.1053 | 10 | -246.0462 | 0.45595
19 V, 50 KHz |  | 18.9907 | 130.53 | 18.9975 | 19.0025 | -487.8947 | 20 | -369.468 | 1.2838
19 V, 100 KHz |  | 18.9855 | 251.05 | 18.9952 | 19.0048 | -762.1053 | 30 | -301.4226 | 1.9373
19 V, 200 KHz |  | 18.9747 | 589.47 | 18.9887 | 19.0113 | -1333.1579 | 70 | -224.5842 | 1.3944
19 V, 300 KHz |  | 18.967 | 589.47 | 18.9887 | 19.0113 | -1737.3684 | 70 | -292.6776 | 1.4385
19 V, 500 KHz |  | 18.98 | 1463.2 | 18.9721 | 19.0279 | -1051.0526 | 100 | -71.6653 | 1.2755
19 V, 1 MHz |  | 19.202 | 3073.7 | 18.9415 | 19.0585 | 10630.5263 | 200 | 345.1246 | 10.4643
100 V, 10 Hz |  | 99.7372 | 600 | 99.94 | 100.06 | -2628 | 10 | -437.9392 | 12.4738
100 V, 20 Hz |  | 99.9305 | 180 | 99.982 | 100.018 | -695 | 10 | -385.5166 | 3.538
100 V, 30 Hz |  | 99.9675 | 180 | 99.982 | 100.018 | -325 | 10 | -180.2776 | 2.0007
100 V, 40 Hz |  | 99.9801 | 85 | 99.9914 | 100.0086 | -199 | 10 | -232.5141 | 3.0191
100 V, 55 Hz |  | 99.988 | 85 | 99.9914 | 100.0086 | -120 | 10 | -140.2095 | 1.5813
100 V, 300 Hz |  | 99.997 | 85 | 99.9914 | 100.0086 | -30 | 10 | -35.0524 | 1.9003
100 V, 1 KHz |  | 99.9974 | 85 | 99.9914 | 100.0086 | -26 | 10 | -30.3787 | 1.1667
100 V, 10 KHz |  | 99.9867 | 85 | 99.9914 | 100.0086 | -133 | 10 | -155.3989 | 0.97195
100 V, 20 KHz |  | 99.9788 | 85 | 99.9914 | 100.0086 | -212 | 10 | -247.7034 | 2.4726
100 V, 30 KHz |  | 99.9701 | 240 | 99.976 | 100.024 | -299 | 10 | -124.4753 | 1.4244
100 V, 50 KHz |  | 99.9564 | 240 | 99.9759 | 100.0241 | -436 | 20 | -181.0391 | 2.5232
100 V, 100 KHz |  | 99.921 | 600 | 99.9399 | 100.0601 | -790 | 30 | -131.5024 | 5.0894
100 V, 200 KHz |  | 99.8343 | 2600 | 99.74 | 100.26 | -1657 | 50 | -63.719 | 14.6283
1000 V, 55 Hz |  | 999.819 | 79 | 999.9196 | 1000.0804 | -181 | 15 | -225.0923 | 1.0002
1000 V, 300 Hz |  | 999.933 | 79 | 999.9196 | 1000.0804 | -67 | 15 | -83.3215 | 1.5001
700 V, 1 KHz |  | 699.968 | 80.714 | 699.9335 | 700.0665 | -45.7143 | 50 | -48.1477 | 1.9048
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.9895e-07 | 10055.5583 |  |  | -1050 |  |  | 95.8433
10 µADC | 1e-05 | 9.9993e-06 | 1045.0732 |  |  | -70 |  |  | 5.0004
50 µADC | 5e-05 | 5.0001e-05 | 245 |  |  | 20 |  |  | 0.99998
100 µADC | 0.0001 | 0.0001 | 145 | 9.9985e-05 | 0.00010001 | 31 | 7 | 21.3544 | 0.49998
-100 µADC | -0.0001 | -0.00010001 | 145 | -9.9985e-05 | -0.00010001 | 63.5 | 7 | 43.7422 | -0.57731
-50 µADC | -5e-05 | -5.0004e-05 | 245 |  |  | 72 |  |  | -0.99993
-10 µADC | -1e-05 | -1.0002e-05 | -1044.8433 |  |  | 150 |  |  | -4.9993
0.1 mADC | 0.0001 | 9.9998e-05 | 145 |  |  | -15 |  |  | 5.7736
0.5 mADC | 0.0005 | 0.00050002 | 65 |  |  | 30 |  |  | 0
1 mADC | 0.001 | 0.001 | 55 | 0.00099994 | 0.0010001 | 39 | 7 | 70.3417 | 0.49998
-1 mADC | -0.001 | -0.0010001 | 55 | -0.00099994 | -0.0010001 | 62 | 7 | 111.8252 | 0
-0.5 mADC | -0.0005 | -0.00050003 | 65 |  |  | 68 |  |  | -0.99993
1 mADC | 0.001 | 0.001 | 55 |  |  | -5 |  |  | 5.7735
5 mADC | 0.005 | 0.0050002 | 65 |  |  | 38 |  |  | 0.99996
10 mADC | 0.01 | 0.01 | 55 | 0.0099994 | 0.010001 | 45 | 7 | 81.1635 | 0.49998
-10 mADC | -0.01 | -0.010001 | 55 | -0.0099994 | -0.010001 | 62 | 7 | 111.8252 | -0.49997
-5 mADC | -0.005 | -0.0050003 | 65 |  |  | 66 |  |  | 0
10 mADC | 0.01 | 0.0099997 | 55 |  |  | -30 |  |  | 5.0002
50 mADC | 0.05 | 0.050002 | 75 |  |  | 44 |  | 58.6667 | 0.99996
100 mADC | 0.1 | 0.10001 | 65 | 0.099993 | 0.10001 | 50.5 | 7 | 77.2457 | 0.57732
-100 mADC | -0.1 | -0.10001 | 65 | -0.099993 | -0.10001 | 70 | 7 | 107.0732 | 0
-50 mADC | -0.05 | -0.050004 | 75 |  |  | 76 |  |  | -0.99992
0.1 ADC | 0.1 | 0.1 | 65 |  |  | 0 |  |  | 5
0.5 ADC | 0.5 | 0.5 | 135 |  |  | 4 |  |  | 1
-0.5 ADC | -0.5 | -0.50004 | 135 |  |  | 78 |  |  | -1.6329
-1.0 ADC | -1 | -1.0001 | 105 | -0.99989 | -1.0001 | 64 | 20 | 59.8759 | -0.49997
1.0 ADC | 1 | 0.99999 | 105 | 0.99989 | 1.0001 | -9 | 20 | -8.42 | 0.5
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.977e-06 | 3658.4144 |  |  | -2300 |  |  | 0
10 µA, 20 Hz | 1e-05 | 9.997e-06 | 2850.8553 |  |  | -300 |  |  | 50.015
10 µA, 30 Hz | 1e-05 | 1e-05 | 2850 |  |  | 0 |  |  | 0
10 µA, 40 Hz | 1e-05 | 1.0001e-05 | 2119.788 |  |  | 100 |  |  | 81.6415
10 µA 55 Hz | 1e-05 | 1.0002e-05 | 2119.4701 |  |  | 250 |  |  | 57.7206
10 µA, 300 Hz | 1e-05 | 1.0004e-05 | 2119.1523 |  |  | -40 |  |  | 49.98
10 µA, 1 KHz | 1e-05 | 1.0003e-05 | 2119.3642 |  |  | 300 |  |  | 81.6252
10 µA, 5 KHz | 1e-05 | 9.9965e-06 | 5501.9257 |  |  | -350 |  |  | 170.8423
10 µA, 10 KHz | 1e-05 | 9.9765e-06 | 11527.0887 |  |  | -2350 |  |  | 57.871
100 µA, 10 Hz | 0.0001 | 9.9779e-05 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 | -4.2047 | 35.0777
100 µA, 20 Hz | 0.0001 | 9.9964e-05 | 600 | 9.994e-05 | 0.00010006 | -360 | 50 | -59.7927 | 0
100 µA, 30 Hz | 0.0001 | 0.0001 | 600 | 9.9939e-05 | 0.00010006 | 0 | 50 | 0 | 0
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 | 9.9968e-05 | 0.00010003 | 130 | 50 | 40.138 | 4.9994
100 µA 55 Hz | 0.0001 | 0.00010002 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 | -12.3502 | 5.7723
100 µA, 300 Hz | 0.0001 | 0.00010003 | 320 | 9.9968e-05 | 0.00010003 | 285 | 50 | 87.9948 | 5.7719
100 µA, 1 KHz | 0.0001 | 0.00010003 | 320 | 9.9939e-05 | 0.00010006 | 250 | 50 | 77.1884 | 8.1629
100 µA, 5 KHz | 0.0001 | 9.997e-05 | 0.1 | 9.999e-05 | 0.00010001 | -300 | 100 | -299.9998 | 16.3348
100 µA, 10 KHz | 0.0001 | 9.9797e-05 | 0.25 | 9.9939e-05 | 0.00010006 | -2030 | 300 | -676.6664 | 8.1816
1mA, 10 Hz | 0.001 | 0.00099781 | 700 | 0.0009993 | 0.0010007 | -2195 | 40 | -313.0607 | 40.9146
1 mA, 20 Hz | 0.001 | 0.00099962 | 390 | 9.9939e-05 | 0.00010006 | -380 | 40 | -96.9274 | 5.0019
1 mA, 30 Hz | 0.001 | 0.00099999 | 390 | 0.00099961 | 0.0010004 | -10 | 40 | -2.5507 | 5.0001
1 mA, 40 Hz | 0.001 | 0.0010001 | 160 | 9.9939e-05 | 0.00010006 | 115 | 40 | 69.729 | 5.7728
1 mA, 55 Hz | 0.001 | 0.0010002 | 160 | 0.00099984 | 0.0010002 | 190 | 40 | 115.2044 | 12.5807
1 mA, 300 Hz | 0.001 | 0.0010003 | 160 | 9.9939e-05 | 0.00010006 | 270 | 40 | 163.7115 | 17.3158
1 mA, 1 KHz | 0.001 | 0.0010003 | 160 | 0.00099984 | 0.0010002 | 250 | 40 | 151.5848 | 8.1629
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 | 30 | 70 | 2.9927 | 8.1647
1 mA, 10 KHz | 0.001 | 0.00099937 | 0.25 | 0.0009998 | 0.0010002 | -630 | 200 | -314.9998 | 15.0095
10 mA, 10 Hz | 0.01 | 0.0099778 | 700 | 9.9939e-05 | 0.00010006 | -2225 | 40 | -317.3395 | 35.1972
10 mA, 20 Hz | 0.01 | 0.0099962 | 390 | 0.0099961 | 0.010004 | -385 | 40 | -98.2028 | 5.7757
10 mA, 30 Hz | 0.01 | 0.0099998 | 390 | 9.9939e-05 | 0.00010006 | -20 | 40 | -5.1014 | 0
10 mA, 40 Hz | 0.01 | 0.010001 | 160 | 0.0099984 | 0.010002 | 100 | 40 | 60.6339 | 8.1641
10 mA, 55 Hz | 0.01 | 0.010002 | 160 | 9.9939e-05 | 0.00010006 | 180 | 40 | 109.141 | 0
10 mA, 300 Hz | 0.01 | 0.010002 | 160 | 0.0099984 | 0.010002 | 240 | 40 | 145.5214 | 12.58
10 mA, 1 KHz | 0.01 | 0.010002 | 160 | 9.9939e-05 | 0.00010006 | 240 | 40 | 145.5214 | 4.9988
10 mA, 5 KHz | 0.01 | 0.01 | 1000 | 0.00999 | 0.01001 | 25 | 70 | 2.4939 | 5.7734
10 mA, 10 KHz | 0.01 | 0.0099942 | 0.25 | 9.9939e-05 | 0.00010006 | -585 | 200 | -292.4998 | 9.5799
100 mA, 10 Hz | 0.1 | 0.09978 | 700 | 0.09993 | 0.10007 | -2195 | 40 | -313.0607 | 14.1732
100 mA, 20 Hz | 0.1 | 0.099964 | 390 | 9.9939e-05 | 0.00010006 | -360 | 40 | -91.826 | 0
100 mA, 30 Hz | 0.1 | 0.1 | 390 | 0.099961 | 0.10004 | 0 | 40 | 0 | 0
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 | 120 | 40 | 72.7607 | 4.9994
100 mA, 55 Hz | 0.1 | 0.10002 | 160 | 0.099984 | 0.10002 | 200 | 40 | 121.2678 | 0
100 mA, 300 Hz | 0.1 | 0.10003 | 160 | 9.9939e-05 | 0.00010006 | 285 | 40 | 172.8066 | 17.0734
100 mA, 1 KHz | 0.1 | 0.10002 | 160 | 0.099984 | 0.10002 | 220 | 40 | 133.3946 | 11.5445
100 mA, 5 KHz | 0.1 | 0.10002 | 0.1 | 9.9939e-05 | 0.00010006 | 160 | 70 | 228.5712 | 4.9992
100 mA, 10 KHz | 0.1 | 0.099978 | 0.25 | 0.09998 | 0.10002 | -220 | 200 | -109.9999 | 8.1668
1 A, 10 Hz | 1 | 0.99751 | -1 | 9.9939e-05 | 0.00010006 | -2495 | 40 | -6235.5517 | 23.8643
1 A, 20 Hz | 1 | 0.99937 | -1 | 0.99996 | 1 | -630 | 40 | -1574.508 | 5.0032
1 A, 30 Hz | 1 | 0.99974 | -1 | 9.9939e-05 | 0.00010006 | -260 | 40 | -649.797 | 5.0013
1 A, 40 Hz | 1 | 0.99986 | 640 | 0.99936 | 1.0006 | -140 | 40 | -21.8324 | 0
1 A, 55 Hz | 1 | 0.99996 | 640 | 9.9939e-05 | 0.00010006 | -40 | 40 | -6.2378 | 5.0002
1 A, 300 Hz | 1 | 1 | 640 | 0.99936 | 1.0006 | 30 | 40 | 4.6784 | 16.3294
1 A, 1 KHz | 1 | 0.99974 | 640 | 9.9939e-05 | 0.00010006 | -260 | 40 | -40.5459 | 0
1 A, 5 KHz | 1 | 0.99955 | 800 | 0.9992 | 1.0008 | -455 | 70 | -56.6585 | 12.9158
1 A, 10 KHz | 1 | 0.99808 | 0.82 | 9.9939e-05 | 0.00010006 | -1920 | 200 | -959.9919 | 12.6073
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.01 | 277.45 | 291.94 | 66.65 | 20.28 | 23.3 |  | 679.63