## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 43268
Model Number |  | 1281 |  |  | Test date |  | 24-11-01 22:29:13
Meter Info |  |  |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1,1,0,0,0 |  |  | CERT_AMB?
Firmware |  | 890144-03.14 |  |  | TEMP?
Date of Baseline Cal |  |  |  |  | TEMP? BASE
Date of Certified Cal |  |  |  |  | TEMP? CERTIFIED
CNFTST? |  | 0 |  |  | STANDBY?
DUT Confidence level
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 191024 |  |  | MFC last calibrated |  | 13 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 191024 |  |  | MFC since DCV Zero |  | 13 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529912686322891E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306135899178203E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039668828305E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000065537234996E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996275049758838E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6070 days 10.47 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 8.10 hours |  |  | CAL CONST,DACLIN |  | 0
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.083 | 7.5e-07 |  |  | N/A |  |  | 0.033276 | 0.116
0.0VDC | 0 | -1.05 | 7.5e-07 |  |  | N/A |  |  | 0.034681 | -0.99
0.0VDC | 0 | -10.4 | 7.5e-07 |  |  | N/A |  |  | 0.25712 | -10.6
0.0VDC | 0 | -94 | 7.5e-07 |  |  | N/A |  |  | 3.4561 | -92
0.0VDC | 0 | -860 | 7.5e-07 |  |  | N/A |  |  | 23.5112 | -870
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 2.4328 | 0.01
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.54281 | 0.02
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.19936 | 0.05
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.20953 | 0.1
-0.01 | 100mV | -0.0099998 | -81.5 | [5 empty] | -1.7281 | -0.0099998
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.64378 | -0.02
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.1236 | -0.05
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.11203 | -0.1
0.1 | 1V | 0.099998 | 14 | [5 empty] | 0.44002 | 0.099998
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.11024 | 0.2
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.07356 | 0.5
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.053489 | 1
-0.1 | 1V | -0.1 | 14 | [5 empty] | -0.48477 | -0.1
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.082074 | -0.2
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.1253 | -0.5
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.034075 | -1
1 | 10V | 0.99999 | 4.7 | [5 empty] | 0.47112 | 0.99999
2 | 10V | 2 | 4.1 | [5 empty] | 0.082074 | 2
5 | 10V | 5 | 4.1 | [5 empty] | 0.070554 | 5
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.01236 | 10
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.020572 | 19
-1 | 10V | -1 | 4.7 | [5 empty] | -0.7211 | -1
-2 | 10V | -2 | 4.1 | [5 empty] | -0.14093 | -2
-5 | 10V | -5 | 4.1 | [5 empty] | -0.084261 | -5
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.018105 | -10
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.031408 | -19
10 | 100V | 9.9999 | 3.8 | [5 empty] | 0.79374 | 9.9999
20 | 100V | 19.9998 | 3.8 | [5 empty] | 0.16976 | 19.9998
50 | 100V | 49.9997 | 7 | [5 empty] | 0.10573 | 49.9997
100 | 100V | 99.9994 | 6 |  |  |  | 2 |  | 0.018559 | 99.9994
-10 | 100V | -10.0001 | 3.8 | [5 empty] | -0.50935 | -10
-20 | 100V | -20 | 3.8 | [5 empty] | -0.11667 | -20
-50 | 100V | -49.9999 | 7 | [5 empty] | -0.149 | -49.9999
-100 | 100V | -99.9996 | 6 |  |  |  | 2 |  | -0.075185 | -99.9996
-100 | 1000V | -100.0005 | 6 | [5 empty] | -0.4272 | -100.0006
-200 | 1000V | -200.0001 | 5.5 | [5 empty] | -0.14142 | -200.0001
-500 | 1000V | -499.9989 | 8.2 | [5 empty] | -0.060369 | -499.9989
-1000 | 1000V | -999.9953 | 7.6 |  |  |  | 2 |  | -0.06766 | -999.9952
100 | 1000V | 99.9987 | 6 | [5 empty] | 0.40139 | 99.9987
200 | 1000V | 199.9982 | 5.5 | [5 empty] | 0.125 | 199.9983
500 | 1000V | 499.997 | 8.2 | [5 empty] | 0.060919 | 499.9971
1000 | 1000V | 999.9938 | 7.6 |  |  |  | 2 |  | 0.06225 | 999.9936
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99963 | 0.9997 | 85 |  |  |  | 20 |  | 1.1779e-10 | 0.9997
1.9 Ω | 1.9 | 1.9 | 85 |  |  |  | 15 |  | 1.2395e-10 | 1.9
10 Ω | 10.0003 | 10.0004 | 26 |  |  |  | 5 |  | 0 | 10.0004
19 Ω | 19.0038 | 19.0039 | 24 |  |  |  | 5 |  | 1.754 | 19.0039
100 Ω | 100.0049 | 100.0047 | 15 |  |  |  | 3 |  | 0.52702 | 100.0047
190 Ω | 190.0023 | 190.0025 | 15 |  |  |  | 3 |  | 0 | 190.0024
1 kΩ | 999.9418 | 999.938 | 11 |  |  |  | 3 |  | 0.52708 | 999.938
1.9 kΩ | 1899.9019 | 1899.898 | 11 |  |  |  | 3 |  | 0.2321 | 1899.897
10 kΩ | 10000.397 | 10000.38 | 9 |  |  |  | 3 |  | 0.52703 | 10000.38
19 kΩ | 18999.633 | 18999.59 | 9 |  |  |  | 3 |  | 0.46418 | 18999.59
100 kΩ | 99998.22 | 99998 | 11 |  |  |  | 5 |  | 0 | 99998
190 kΩ | 189993.74 | 189993.3 | 11 |  |  |  | 5 |  | 0.55481 | 189993.2
1 MΩ | 999904.9 | 999900 | 16 |  |  |  | 8 |  | 2.1084 | 999903
1.9 MΩ | 1899848.4 | 1899840 | 17 |  |  |  | 8 |  | 0.58849 | 1899840
10MΩ | 9998508 | 9998370 | 33 |  |  |  | 12 |  | 6.5352 | 9998480
19 MΩ | 18999402 | 18998550 | 43 |  |  |  | 12 |  | 2.8763 | 18998510
100MΩ(2W) | 100009510 | 2e+35 | 110 |  |  |  | 180 |  | 0 | 2e+35
| 100009510 | 2e+35 | 110 | [5 empty] | 0 | 2e+35
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0 | 5e-05 |  |  | N/A |  |  | 5.2705e-05 | -0.0001
0 Ω | 100 Ω | 0 | 5e-05 |  |  | N/A |  |  | 0 | 0
0 Ω | 1 KΩ | -0.0001 | 5e-05 |  |  | N/A |  |  | 4.4096e-05 | 0
0 Ω | 10 KΩ | -0.001 | 5e-05 |  |  | N/A |  |  | 0.0005 | -0.001
0 Ω | 100 KΩ | -0.01 | 5e-05 |  |  | N/A |  |  | 0.0033333 | -0.02
0 Ω | 1 MΩ | -0.1 | 5e-05 |  |  | N/A |  |  | 0 | -0.1
0 Ω | 10 MΩ | -1 | 5e-05 |  |  | N/A |  |  | 0.33333 | -1
0 Ω | 100 MΩ(2W) | -10 | 5e-05 |  |  | N/A |  |  | 0 | -10
|  | -10 | 5e-05 | [5 empty] | 0 | -10
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 1.3873 | 5e-05 |  |  | N/A |  |  | 0 | 1.3872
0 Ω | 100 Ω | 1.3872 | 5e-05 |  |  | N/A |  |  | 4.4096e-05 | 1.3872
0 Ω | 1 KΩ | 1.3872 | 5e-05 |  |  | N/A |  |  | 5.2705e-05 | 1.3871
0 Ω | 10 KΩ | 1.386 | 5e-05 |  |  | N/A |  |  | 0.00033333 | 1.386
0 Ω | 100 KΩ | 1.37 | 5e-05 |  |  | N/A |  |  | 0.0044096 | 1.38
0 Ω | 1 MΩ | 1.3 | 5e-05 |  |  | N/A |  |  | 0 | 1.3
0 Ω | 10 MΩ | 1 | 5e-05 |  |  | N/A |  |  | 0 | 1
0 Ω | 100 MΩ | -10 | 5e-05 |  |  | N/A |  |  | 4.4096 | -10
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 1 | 77 |  |  |  | 10 |  | 13.3331 | 1
2VAC/1kHz | 10 VAC | 2 | 73.5 |  |  |  | 10 |  | 2.6352 | 2
5VAC/1kHz | 10 VAC | 5.0001 | 84 |  |  |  | 10 |  | 2.1858 | 5.0001
10VAC/1kHz | 10 VAC | 10.0002 | 77 |  |  |  | 10 |  | 1.7873 | 10.0002
12VAC/1kHz | 10 VAC | 12.0002 | 75.833 |  |  |  | 10 |  | 1.1866 | 12.0002
15VAC/1kHz | 10 VAC | 15.0002 | 74.667 |  |  |  | 10 |  | 1.0061 | 15.0002
19VAC/1kHz | 10 VAC | 19.0002 | 73.684 |  |  |  | 10 |  | 1.4171 | 19.0002
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0010802 | 5500 |  |  |  | 20 |  | 37742.8483 | 0.0011251
0.001 V, 20 Hz |  | 0.0009951 | 5200 |  |  |  | 20 |  | 6253.27 | 0.0009885
0.001 V, 30 Hz |  | 0.0010001 | 5200 |  |  |  | 20 |  | 1213.8025 | 0.0009988
0.001 V, 40 Hz |  | 0.0010032 | 5100 |  |  |  | 20 |  | 725.6888 | 0.0010039
0.001 V, 55 Hz |  | 0.0010024 | 5100 |  |  |  | 20 |  | 1527.7633 | 0.0010011
0.001 V, 300 Hz |  | 0.0010011 | 5100 |  |  |  | 20 |  | 843.1697 | 0.0010003
0.001 V, 1 KHz |  | 0.001003 | 5100 |  |  |  | 20 |  | 566.1923 | 0.0010039
0.001 V, 10 KHz |  | 0.0010041 | 5100 |  |  |  | 20 |  | 954.9613 | 0.0010036
0.001 V, 20 KHz |  | 0.0010062 | 5100 |  |  |  | 20 |  | 551.3588 | 0.0010066
0.001 V, 30 KHz |  | 0.0010071 | 5340 |  |  |  | 20 |  | 437.8498 | 0.0010081
0.001 V, 50 KHz |  | 0.0010056 | 5340 |  |  |  | 20 |  | 322.2325 | 0.0010052
0.001 V, 100 KHz |  | 0.0009812 | 8800 |  |  |  | 50 |  | 297.1337 | 0.0009812
0.001 V, 200 KHz |  | 0.0009052 | 16100 |  |  |  | 100 |  | 207.4942 | 0.0009051
0.001 V, 300 KHz |  | 0.0008644 | 16100 |  |  |  | 100 |  | 263.6664 | 0.000865
0.001 V, 500 KHz |  | 0.0008174 | 31500 |  |  |  | 200 |  | 237.7845 | 0.0008174
0.001 V, 1 MHz |  | 0.0006558 | 33000 |  |  |  | 300 |  | 366.5296 | 0.0006558
0.01 V, 10 Hz |  | 0.010839 | 1100 |  |  |  | 20 |  | 32965.6905 | 0.011138
0.01 V, 20 Hz |  | 0.0099508 | 800 |  |  |  | 20 |  | 1785.1965 | 0.0099611
0.01 V, 30 Hz |  | 0.0099776 | 800 |  |  |  | 20 |  | 296.6374 | 0.0099769
0.01 V, 40 Hz |  | 0.0099949 | 700 |  |  |  | 20 |  | 90.8759 | 0.0099937
0.01 V, 55 Hz |  | 0.0099987 | 700 |  |  |  | 20 |  | 84.8638 | 0.01
0.01 V, 300 Hz |  | 0.0099988 | 700 |  |  |  | 20 |  | 68.4025 | 0.0099988
0.01 V, 1 KHz |  | 0.0099991 | 700 |  |  |  | 20 |  | 86.0471 | 0.0099993
0.01 V, 10 KHz |  | 0.010006 | 700 |  |  |  | 20 |  | 54.9939 | 0.010005
0.01 V, 20 KHz |  | 0.010017 | 700 |  |  |  | 20 |  | 50.9288 | 0.010017
0.01 V, 30 KHz |  | 0.010025 | 940 |  |  |  | 20 |  | 47.433 | 0.010025
0.01 V, 50 KHz |  | 0.010025 | 940 |  |  |  | 20 |  | 22.306 | 0.010024
0.01 V, 100 KHz |  | 0.0099744 | 1600 |  |  |  | 50 |  | 31.3051 | 0.0099743
0.01 V, 200 KHz |  | 0.0098497 | 2600 |  |  |  | 100 |  | 18.9938 | 0.0098498
0.01 V, 300 KHz |  | 0.0097924 | 2600 |  |  |  | 100 |  | 16.5015 | 0.0097921
0.01 V, 500 KHz |  | 0.0097345 | 4500 |  |  |  | 200 |  | 29.4067 | 0.0097346
0.01 V, 1 MHz |  | 0.0088285 | 6000 |  |  |  | 300 |  | 36.46 | 0.0088283
0.1 V, 10 Hz |  | 0.10833 | 660 |  |  |  | 20 |  | 32247.3069 | 0.10851
0.1 V, 20 Hz |  | 0.099621 | 300 |  |  |  | 20 |  | 1698.7419 | 0.099614
0.1 V, 30 Hz |  | 0.099793 | 300 |  |  |  | 20 |  | 222.8449 | 0.099787
0.1 V, 40 Hz |  | 0.099943 | 195 |  |  |  | 20 |  | 94.2719 | 0.099949
0.1 V, 55 Hz |  | 0.099994 | 195 |  |  |  | 20 |  | 16.4208 | 0.099992
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 4.9497 | 0.1
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 5.979 | 0.1
0.1 V, 10 KHz |  | 0.10005 | 195 |  |  |  | 20 |  | 8.1628 | 0.10005
0.1 V, 20 KHz |  | 0.10014 | 195 |  |  |  | 20 |  | 6.9384 | 0.10014
0.1 V, 30 KHz |  | 0.1002 | 400 |  |  |  | 20 |  | 2.3989 | 0.1002
0.1 V, 50 KHz |  | 0.10014 | 400 |  |  |  | 20 |  | 3.5461 | 0.10014
0.1 V, 100 KHz |  | 0.099507 | 1050 |  |  |  | 50 |  | 2.1966 | 0.099506
0.1 V, 200 KHz |  | 0.098139 | 1240 |  |  |  | 100 |  | 2.6364 | 0.098139
0.1 V, 300 KHz |  | 0.097305 | 1240 |  |  |  | 100 |  | 2.8352 | 0.097306
0.1 V, 500 KHz |  | 0.095649 | 1900 |  |  |  | 200 |  | 5.7975 | 0.09565
0.1 V, 1 MHz |  | 0.084648 | 4000 |  |  |  | 300 |  | 18.7071 | 0.084648
1 V, 10 Hz |  | 1.0878 | 600 |  |  |  | 10 |  | 31794.8737 | 1.0451
1 V, 20 Hz |  | 0.99641 | 180 |  |  |  | 10 |  | 1620.0876 | 0.99779
1 V, 30 Hz |  | 0.998 | 180 |  |  |  | 10 |  | 231.0159 | 0.99824
1 V, 40 Hz |  | 0.99952 | 77 |  |  |  | 10 |  | 71.9916 | 0.99958
1 V, 55 Hz |  | 0.99993 | 77 |  |  |  | 10 |  | 4.8937 | 0.99993
1 V, 300 Hz |  | 1 | 77 |  |  |  | 10 |  | 3.2318 | 1
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.5634 | 1
1 V, 10 KHz |  | 1.0004 | 77 |  |  |  | 10 |  | 0.97143 | 1.0004
1 V, 20 KHz |  | 1.0012 | 77 |  |  |  | 10 |  | 1.9976 | 1.0012
1 V, 30 KHz |  | 1.0016 | 140 |  |  |  | 10 |  | 0.92646 | 1.0016
1 V, 50 KHz |  | 1.001 | 140 |  |  |  | 20 |  | 1.801 | 1.001
1 V, 100 KHz |  | 0.9947 | 310 |  |  |  | 30 |  | 2.7532 | 0.99471
1 V, 200 KHz |  | 0.9815 | 550 |  |  |  | 70 |  | 1.074 | 0.9815
1 V, 300 KHz |  | 0.97415 | 550 |  |  |  | 70 |  | 2.7534 | 0.97415
1 V, 500 KHz |  | 0.96922 | 1400 |  |  |  | 100 |  | 2.0707 | 0.96921
1 V, 1 MHz |  | 0.97065 | 3000 |  |  |  | 200 |  | 3.5315 | 0.97066
10 V, 10 Hz |  | 10.775 | 600 |  |  |  | 10 |  | 29811.4948 | 10.9472
10 V, 20 Hz |  | 9.965 | 180 |  |  |  | 10 |  | 1777.1745 | 9.9523
10 V, 30 Hz |  | 9.9797 | 180 |  |  |  | 10 |  | 147.8006 | 9.9813
10 V, 40 Hz |  | 9.9953 | 77 |  |  |  | 10 |  | 68.7691 | 9.9961
10 V, 55 Hz |  | 9.9994 | 77 |  |  |  | 10 |  | 7.4671 | 9.9993
10 V, 300 Hz |  | 10.0001 | 77 |  |  |  | 10 |  | 2.1213 | 10.0001
10 V, 1 KHz |  | 10.0002 | 77 |  |  |  | 10 |  | 2.1213 | 10.0001
10 V, 10 KHz |  | 10.0033 | 77 |  |  |  | 10 |  | 1.3638 | 10.0032
10 V, 20 KHz |  | 10.0099 | 77 |  |  |  | 10 |  | 0.97086 | 10.0099
10 V, 30 KHz |  | 10.0136 | 140 |  |  |  | 10 |  | 2.1053 | 10.0136
10 V, 50 KHz |  | 10.0068 | 140 |  |  |  | 20 |  | 1.9208 | 10.0068
10 V, 100 KHz |  | 9.9461 | 270 |  |  |  | 30 |  | 1.7495 | 9.9461
10 V, 200 KHz |  | 9.8191 | 670 |  |  |  | 70 |  | 1.3472 | 9.8191
10 V, 300 KHz |  | 9.7498 | 670 |  |  |  | 70 |  | 1.5194 | 9.7498
10 V, 500 KHz |  | 9.7127 | 1700 |  |  |  | 100 |  | 2.2109 | 9.7126
10 V, 1 MHz |  | 9.819 | 3500 |  |  |  | 200 |  | 6.674 | 9.8189
19 V, 10 Hz |  | 2e+35 | 552.63 |  |  |  | 10 |  | 0 | 19.8752
19 V, 20 Hz |  | 18.9212 | 165.79 |  |  |  | 10 |  | 1601.4522 | 18.9585
19 V, 30 Hz |  | 18.9606 | 165.79 |  |  |  | 10 |  | 220.5727 | 18.9602
19 V, 40 Hz |  | 18.9913 | 73.684 |  |  |  | 10 |  | 65.6473 | 18.9908
19 V, 55 Hz |  | 18.9988 | 73.684 |  |  |  | 10 |  | 4.3324 | 18.9987
19 V, 300 Hz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 2.8097 | 19.0002
19 V, 1 KHz |  | 19.0002 | 73.684 |  |  |  | 10 |  | 1.6317 | 19.0002
19 V, 10 KHz |  | 19.0059 | 73.684 |  |  |  | 10 |  | 0.98824 | 19.0059
19 V, 20 KHz |  | 19.0185 | 73.684 |  |  |  | 10 |  | 1.1154 | 19.0185
19 V, 30 KHz |  | 19.0253 | 130.53 |  |  |  | 10 |  | 0.98334 | 19.0253
19 V, 50 KHz |  | 19.012 | 130.53 |  |  |  | 20 |  | 3.5295 | 19.012
19 V, 100 KHz |  | 18.8954 | 251.05 |  |  |  | 30 |  | 1.543 | 18.8953
19 V, 200 KHz |  | 18.6501 | 589.47 |  |  |  | 70 |  | 1.3523 | 18.6501
19 V, 300 KHz |  | 18.5142 | 589.47 |  |  |  | 70 |  | 0.76914 | 18.5142
19 V, 500 KHz |  | 18.4288 | 1463.2 |  |  |  | 100 |  | 0.89071 | 18.4288
19 V, 1 MHz |  | 18.5728 | 3073.7 |  |  |  | 200 |  | 8.6012 | 18.573
100 V, 10 Hz |  | 105.687 | 600 |  |  |  | 10 |  | 34396.681 | 105.1681
100 V, 20 Hz |  | 99.6292 | 180 |  |  |  | 10 |  | 1829.9387 | 99.6253
100 V, 30 Hz |  | 99.8123 | 180 |  |  |  | 10 |  | 223.2934 | 99.8071
100 V, 40 Hz |  | 99.957 | 85 |  |  |  | 10 |  | 74.6496 | 99.9468
100 V, 55 Hz |  | 99.9972 | 85 |  |  |  | 10 |  | 6.5851 | 99.9976
100 V, 300 Hz |  | 100.005 | 85 |  |  |  | 10 |  | 1.74 | 100.0047
100 V, 1 KHz |  | 100.006 | 85 |  |  |  | 10 |  | 1.9435 | 100.0064
100 V, 10 KHz |  | 100.0425 | 85 |  |  |  | 10 |  | 1.3223 | 100.0423
100 V, 20 KHz |  | 100.1114 | 85 |  |  |  | 10 |  | 0.72567 | 100.1113
100 V, 30 KHz |  | 100.1461 | 240 |  |  |  | 10 |  | 1.2674 | 100.1462
100 V, 50 KHz |  | 100.0712 | 240 |  |  |  | 20 |  | 2.0055 | 100.0716
100 V, 100 KHz |  | 99.4549 | 600 |  |  |  | 30 |  | 6.1778 | 99.4548
100 V, 200 KHz |  | 98.1774 | 2600 |  |  |  | 50 |  | 11.8029 | 98.1757
1000 V, 55 Hz |  | 1009.872 | 79 |  |  |  | 15 |  | 7.0793 | 1009.867
1000 V, 300 Hz |  | 1009.975 | 79 |  |  |  | 15 |  | 2.2747 | 1009.977
700 V, 1 KHz |  | 709.972 | 80.714 |  |  |  | 50 |  | 1.7878 | 709.971
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 1.0003e-06 | 10041.9874 | [5 empty] | 49.985 | 1.0003e-06
10 µADC | 1e-05 | 1e-05 | 1044.9896 | [5 empty] | 5 | 1e-05
50 µADC | 5e-05 | 5e-05 | 245 | [5 empty] | 1.633 | 5e-05
100 µADC | 0.0001 | 9.9999e-05 | 145 |  |  |  | 7 |  | 0.57735 | 9.9999e-05
-100 µADC | -0.0001 | -9.9999e-05 | 145 |  |  |  | 7 |  | -0.50001 | -9.9999e-05
-50 µADC | -5e-05 | -4.9999e-05 | 245 | [5 empty] | -1 | -4.9999e-05
-10 µADC | -1e-05 | -9.9997e-06 | -1045.0366 | [5 empty] | -9.5746 | -9.9997e-06
0.1 mADC | 0.0001 | 0.00010004 | 145 | [5 empty] | 0 | 0.00010004
0.5 mADC | 0.0005 | 0.00050002 | 65 | [5 empty] | 0 | 0.00050002
1 mADC | 0.001 | 0.00099999 | 55 |  |  |  | 7 |  | 0 | 0.00099999
-1 mADC | -0.001 | -0.00099999 | 55 |  |  |  | 7 |  | -0.5 | -0.00099999
-0.5 mADC | -0.0005 | -0.00049997 | 65 | [5 empty] | -1.1548 | -0.00049997
1 mADC | 0.001 | 0.0010004 | 55 | [5 empty] | 0 | 0.0010004
5 mADC | 0.005 | 0.0050003 | 65 | [5 empty] | 0.99994 | 0.0050003
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0.57735 | 0.01
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | -0.95742 | -0.01
-5 mADC | -0.005 | -0.0049998 | 65 | [5 empty] | 0 | -0.0049998
10 mADC | 0.01 | 0.010004 | 55 | [5 empty] | 4.9979 | 0.010004
50 mADC | 0.05 | 0.050002 | 75 | [5 empty] | 1.1546 | 0.050002
100 mADC | 0.1 | 0.1 | 65 |  |  |  | 7 |  | 0.57735 | 0.1
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.5 | -0.1
-50 mADC | -0.05 | -0.049997 | 75 | [5 empty] | -1.0001 | -0.049998
0.1 ADC | 0.1 | 0.10005 | 65 | [5 empty] | 4.9973 | 0.10005
0.5 ADC | 0.5 | 0.50006 | 135 | [5 empty] | 0 | 0.50006
-0.5 ADC | -0.5 | -0.50001 | 135 | [5 empty] | -0.99998 | -0.50001
-1.0 ADC | -1 | -1 | 105 |  |  |  | 20 |  | -0.8165 | -1
1.0 ADC | 1 | 0.99999 | 105 |  |  |  | 20 |  | 0.95743 | 0.99999
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 1.1399e-05 | 3202.0353 | [5 empty] | 129279.9535 | 9.712e-06
10 µA, 20 Hz | 1e-05 | 9.97e-06 | 2858.5757 | [5 empty] | 5828.6102 | 9.983e-06
10 µA, 30 Hz | 1e-05 | 9.975e-06 | 2857.1429 | [5 empty] | 866.2644 | 9.97e-06
10 µA, 40 Hz | 1e-05 | 1.0006e-05 | 2118.7288 | [5 empty] | 3680.5217 | 9.988e-06
10 µA 55 Hz | 1e-05 | 9.9965e-06 | 2120.7423 | [5 empty] | 2802.171 | 9.997e-06
10 µA, 300 Hz | 1e-05 | 1.0008e-05 | 2118.3054 |  |  | -40 |  |  | 49.96 | 1.0008e-05
10 µA, 1 KHz | 1e-05 | 1.0007e-05 | 2118.517 | [5 empty] | 287.0272 | 1.0007e-05
10 µA, 5 KHz | 1e-05 | 1.0005e-05 | 5497.2514 | [5 empty] | 149.925 | 1.0005e-05
10 µA, 10 KHz | 1e-05 | 9.992e-06 | 11509.2074 | [5 empty] | 0 | 9.992e-06
100 µA, 10 Hz | 0.0001 | 9.8649e-05 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 153378.7003 | 0.00011639
100 µA, 20 Hz | 0.0001 | 9.9794e-05 | 600 |  |  |  | 50 |  | 7206.3527 | 9.8707e-05
100 µA, 30 Hz | 0.0001 | 9.9716e-05 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 248.8149 | 9.9862e-05
100 µA, 40 Hz | 0.0001 | 9.9957e-05 | 320 |  |  |  | 50 |  | 439.6132 | 9.9977e-05
100 µA 55 Hz | 0.0001 | 9.9998e-05 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 232.4386 | 9.9969e-05
100 µA, 300 Hz | 0.0001 | 0.0001 | 320 |  |  |  | 50 |  | 12.5828 | 0.0001
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 20.6151 | 0.0001
100 µA, 5 KHz | 0.0001 | 9.9972e-05 | 0.1 |  |  |  | 100 |  | 9.5769 | 9.9972e-05
100 µA, 10 KHz | 0.0001 | 9.9871e-05 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 5.0065 | 9.987e-05
1mA, 10 Hz | 0.001 | 0.0011566 | 700 |  |  |  | 40 |  | 126544.8385 | 0.0012529
1 mA, 20 Hz | 0.001 | 0.00099436 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 6447.8174 | 0.00099515
1 mA, 30 Hz | 0.001 | 0.00099858 | 390 |  |  |  | 40 |  | 457.4135 | 0.00099686
1 mA, 40 Hz | 0.001 | 0.00099934 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 378.5244 | 0.00099915
1 mA, 55 Hz | 0.001 | 0.00099978 | 160 |  |  |  | 40 |  | 38.7385 | 0.00099975
1 mA, 300 Hz | 0.001 | 0.00099995 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 18.2583 | 0.00099993
1 mA, 1 KHz | 0.001 | 0.00099993 | 160 |  |  |  | 40 |  | 25.1679 | 0.00099994
1 mA, 5 KHz | 0.001 | 0.00099999 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0 | 0.00099997
1 mA, 10 KHz | 0.001 | 0.0010001 | 0.25 |  |  |  | 200 |  | 9.5731 | 0.0010001
10 mA, 10 Hz | 0.01 | 0.011001 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 126369.8536 | 0.012295
10 mA, 20 Hz | 0.01 | 0.009976 | 390 |  |  |  | 40 |  | 7467.3797 | 0.0098744
10 mA, 30 Hz | 0.01 | 0.0099775 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 561.9382 | 0.0099898
10 mA, 40 Hz | 0.01 | 0.0099922 | 160 |  |  |  | 40 |  | 216.8871 | 0.0099896
10 mA, 55 Hz | 0.01 | 0.0099977 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 17.0823 | 0.009998
10 mA, 300 Hz | 0.01 | 0.0099991 | 160 |  |  |  | 40 |  | 17.3221 | 0.0099991
10 mA, 1 KHz | 0.01 | 0.0099992 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 10.0008 | 0.0099993
10 mA, 5 KHz | 0.01 | 0.01 | 1000 |  |  |  | 70 |  | 12.91 | 0.01
10 mA, 10 KHz | 0.01 | 0.010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 4.9989 | 0.010002
100 mA, 10 Hz | 0.1 | 0.098396 | 700 |  |  |  | 40 |  | 147070.2599 | 0.088396
100 mA, 20 Hz | 0.1 | 0.099409 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 7211.2492 | 0.099749
100 mA, 30 Hz | 0.1 | 0.099724 | 390 |  |  |  | 40 |  | 543.4948 | 0.099824
100 mA, 40 Hz | 0.1 | 0.099909 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 261.9488 | 0.099892
100 mA, 55 Hz | 0.1 | 0.099961 | 160 |  |  |  | 40 |  | 9.578 | 0.099965
100 mA, 300 Hz | 0.1 | 0.099974 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 10.0026 | 0.099975
100 mA, 1 KHz | 0.1 | 0.099977 | 160 |  |  |  | 40 |  | 20.6203 | 0.099977
100 mA, 5 KHz | 0.1 | 0.1 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 9.5743 | 0.099998
100 mA, 10 KHz | 0.1 | 0.10005 | 0.25 |  |  |  | 200 |  | 4.9977 | 0.10005
1 A, 10 Hz | 1 | 0.9847 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 150543.7027 | 1.0231
1 A, 20 Hz | 1 | 0.99824 | -1 |  |  |  | 40 |  | 5944.058 | 1.0053
1 A, 30 Hz | 1 | 0.99904 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 370.3461 | 0.99714
1 A, 40 Hz | 1 | 0.99948 | 640 |  |  |  | 40 |  | 237.6642 | 0.99922
1 A, 55 Hz | 1 | 0.99985 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 33.0453 | 0.99989
1 A, 300 Hz | 1 | 0.99998 | 640 |  |  |  | 40 |  | 5.0001 | 0.99995
1 A, 1 KHz | 1 | 0.99992 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 28.725 | 0.99993
1 A, 5 KHz | 1 | 0.99992 | 800 |  |  |  | 70 |  | 9.575 | 0.99992
1 A, 10 KHz | 1 | 0.99954 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 10.0046 | 0.99952
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.01 | 300.97 | 126.42 | 99.74 | 20.5 | 20.61