## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | b'Datron Instruments |  |  | S/N |  | 19346-4
Model Number |  | 1281' |  |  | Test date |  | 25-01-05 20:46:39
Meter Info |  |  |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1,1,0,0,0 |  |  | CERT_AMB?
Firmware |  | 890144-03.12 |  |  | TEMP?
Date of Baseline Cal |  |  |  |  | TEMP? BASE
Date of Certified Cal |  |  |  |  | TEMP? CERTIFIED
CNFTST? |  | 0 |  |  | STANDBY?
DUT Confidence level
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 050125 |  |  | MFC last calibrated |  | 0 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 050125 |  |  | MFC since DCV Zero |  | 0 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529893148158212E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306131483552985E+01 |  |  | CAL CONST, Zero calibration temperature |  | 22
CAL CONST 10KOHM true output resistance |  | +1.000039901931142E+04 |  |  | CAL CONST, ALL calibration temp |  | 22
CAL SONST 10kOHM standard resistance |  | +1.000060285459230E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996120415654161E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6074 days 5.85 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 22.00 hours |  |  | CAL CONST,DACLIN |  | 0
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | -0.1 | 7.5e-07 |  |  | N/A |  |  | 0.009314
0.0VDC | 0 | -0.62 | 7.5e-07 |  |  | N/A |  |  | 0.033166
0.0VDC | 0 | -8 | 7.5e-07 |  |  | N/A |  |  | 0.33082
0.0VDC | 0 | -75 | 7.5e-07 |  |  | N/A |  |  | 2.0276
0.0VDC | 0 | -670 | 7.5e-07 |  |  | N/A |  |  | 25.8736
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.0099999 | 81.5 | [5 empty] | 2.6739
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.29274
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.17256
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.1566
-0.01 | 100mV | -0.01 | -81.5 | [5 empty] | -2.9725
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.6061
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.17901
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.098868
0.1 | 1V | 0.099999 | 14 | [5 empty] | 0.22048
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.17638
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.062539
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.018708
-0.1 | 1V | -0.1 | 14 | [5 empty] | -0.55553
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.1728
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.10817
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.062249
1 | 10V | 0.99999 | 4.7 | [5 empty] | 0.19861
2 | 10V | 2 | 4.1 | [5 empty] | 0.15701
5 | 10V | 5 | 4.1 | [5 empty] | 0.089876
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.042947
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.044513
-1 | 10V | -1 | 4.7 | [5 empty] | -0.48989
-2 | 10V | -2 | 4.1 | [5 empty] | -0.17873
-5 | 10V | -5 | 4.1 | [5 empty] | -0.091348
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.025386
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.019178
10 | 100V | 9.9999 | 3.8 | [5 empty] | 0.35395
20 | 100V | 19.9999 | 3.8 | [5 empty] | 0.16223
50 | 100V | 49.9999 | 7 | [5 empty] | 0.092616
100 | 100V | 99.9999 | 6 |  |  |  | 2 |  | 0.044378
-10 | 100V | -10.0001 | 3.8 | [5 empty] | -0.31623
-20 | 100V | -20.0001 | 3.8 | [5 empty] | -0.10897
-50 | 100V | -50.0001 | 7 | [5 empty] | -0.13055
-100 | 100V | -100 | 6 |  |  |  | 2 |  | -0.055628
-100 | 1000V | -100.0007 | 6 | [5 empty] | -0.2958
-200 | 1000V | -200.0006 | 5.5 | [5 empty] | -0.14458
-500 | 1000V | -500.0011 | 8.2 | [5 empty] | -0.071414
-1000 | 1000V | -999.9999 | 7.6 |  |  |  | 2 |  | -0.02848
100 | 1000V | 99.9994 | 6 | [5 empty] | 0.1453
200 | 1000V | 199.9994 | 5.5 | [5 empty] | 0.15811
500 | 1000V | 499.9999 | 8.2 | [5 empty] | 0.059255
1000 | 1000V | 999.9989 | 7.6 |  |  |  | 2 |  | 0.028333
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99961 | 0.99954 | 85 |  |  |  | 20 |  | 2.7852
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.87285
10 Ω | 10.0003 | 10.0003 | 26 |  |  |  | 5 |  | 0.24888
19 Ω | 19.0038 | 19.0038 | 24 |  |  |  | 5 |  | 0.16313
100 Ω | 100.0049 | 100.005 | 15 |  |  |  | 3 |  | 0.027435
190 Ω | 190.0018 | 190.003 | 15 |  |  |  | 3 |  | 0.020403
1 kΩ | 999.9425 | 999.9417 | 11 |  |  |  | 3 |  | 0.089028
1.9 kΩ | 1899.9023 | 1899.9036 | 11 |  |  |  | 3 |  | 0.046782
10 kΩ | 10000.399 | 10000.4009 | 9 |  |  |  | 3 |  | 0.085988
19 kΩ | 18999.631 | 18999.6386 | 9 |  |  |  | 3 |  | 0.037659
100 kΩ | 99998.21 | 99998.193 | 11 |  |  |  | 5 |  | 0.026823
190 kΩ | 189993.67 | 189993.699 | 11 |  |  |  | 5 |  | 0.010779
1 MΩ | 999904 | 999905.4 | 16 |  |  |  | 8 |  | 0.068522
1.9 MΩ | 1899847.2 | 1899849.09 | 17 |  |  |  | 8 |  | 0.0385
10MΩ | 9998565 | 9998736.8 | 33 |  |  |  | 12 |  | 0.085011
19 MΩ | 18999571 | 18999663.3 | 43 |  |  |  | 12 |  | 0.2544
100MΩ(2W) | 100008360 | 100023200 | 110 |  |  |  | 180 |  | 8.0569
| 100009510 | 99905500 | 110 | [5 empty] | 8.2709
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000106 | 5e-05 |  |  | N/A |  |  | 1.4414e-05
0 Ω | 100 Ω | -6.8e-05 | 5e-05 |  |  | N/A |  |  | 7.7531e-06
0 Ω | 1 KΩ | -5.3e-05 | 5e-05 |  |  | N/A |  |  | 3.9193e-06
0 Ω | 10 KΩ | -0.00067 | 5e-05 |  |  | N/A |  |  | 2.4495e-05
0 Ω | 100 KΩ | -0.0074 | 5e-05 |  |  | N/A |  |  | 0.00036056
0 Ω | 1 MΩ | -0.092 | 5e-05 |  |  | N/A |  |  | 0.0028723
0 Ω | 10 MΩ | -0.83 | 5e-05 |  |  | N/A |  |  | 0.02
0 Ω | 100 MΩ(2W) | -7.9 | 5e-05 |  |  | N/A |  |  | 0.25712
|  | 0 | 5e-05 | [5 empty] | 0
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 1.5058 | 5e-05 |  |  | N/A |  |  | 0.00012047
0 Ω | 100 Ω | 1.5056 | 5e-05 |  |  | N/A |  |  | 7.4978e-05
0 Ω | 1 KΩ | 1.5054 | 5e-05 |  |  | N/A |  |  | 4.0062e-05
0 Ω | 10 KΩ | 1.5044 | 5e-05 |  |  | N/A |  |  | 3.6094e-05
0 Ω | 100 KΩ | 1.4963 | 5e-05 |  |  | N/A |  |  | 0.00043429
0 Ω | 1 MΩ | 1.412 | 5e-05 |  |  | N/A |  |  | 0.0029627
0 Ω | 10 MΩ | 0.66 | 5e-05 |  |  | N/A |  |  | 0.03
0 Ω | 100 MΩ | -6.9 | 5e-05 |  |  | N/A |  |  | 0.36056
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.9999 | 77 |  |  |  | 10 |  | 6.6673
2VAC/1kHz | 10 VAC | 2 | 73.5 |  |  |  | 10 |  | 4.8592
5VAC/1kHz | 10 VAC | 4.9998 | 84 |  |  |  | 10 |  | 2.2362
10VAC/1kHz | 10 VAC | 9.9996 | 77 |  |  |  | 10 |  | 1
12VAC/1kHz | 10 VAC | 11.9995 | 75.833 |  |  |  | 10 |  | 0.84487
15VAC/1kHz | 10 VAC | 14.9993 | 74.667 |  |  |  | 10 |  | 0.67589
19VAC/1kHz | 10 VAC | 18.9991 | 73.684 |  |  |  | 10 |  | 0.43862
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009998 | 5500 |  |  |  | 20 |  | 364.4597
0.001 V, 20 Hz |  | 0.0010013 | 5200 |  |  |  | 20 |  | 412.1115
0.001 V, 30 Hz |  | 0.0010018 | 5200 |  |  |  | 20 |  | 206.4563
0.001 V, 40 Hz |  | 0.0010015 | 5100 |  |  |  | 20 |  | 423.6286
0.001 V, 55 Hz |  | 0.0010013 | 5100 |  |  |  | 20 |  | 556.0536
0.001 V, 300 Hz |  | 0.0010022 | 5100 |  |  |  | 20 |  | 631.9431
0.001 V, 1 KHz |  | 0.0010021 | 5100 |  |  |  | 20 |  | 769.3865
0.001 V, 10 KHz |  | 0.001002 | 5100 |  |  |  | 20 |  | 306.7047
0.001 V, 20 KHz |  | 0.0010029 | 5100 |  |  |  | 20 |  | 612.6336
0.001 V, 30 KHz |  | 0.001003 | 5340 |  |  |  | 20 |  | 366.7012
0.001 V, 50 KHz |  | 0.0010016 | 5340 |  |  |  | 20 |  | 438.0473
0.001 V, 100 KHz |  | 0.0009831 | 8800 |  |  |  | 50 |  | 233.0676
0.001 V, 200 KHz |  | 0.0009204 | 16100 |  |  |  | 100 |  | 242.9452
0.001 V, 300 KHz |  | 0.0008855 | 16100 |  |  |  | 100 |  | 244.6829
0.001 V, 500 KHz |  | 0.0008369 | 31500 |  |  |  | 200 |  | 251.9041
0.001 V, 1 MHz |  | 0.0006493 | 33000 |  |  |  | 300 |  | 405.8573
0.01 V, 10 Hz |  | 0.0099722 | 1100 |  |  |  | 20 |  | 22.8549
0.01 V, 20 Hz |  | 0.0099919 | 800 |  |  |  | 20 |  | 15.0122
0.01 V, 30 Hz |  | 0.0099958 | 800 |  |  |  | 20 |  | 23.1638
0.01 V, 40 Hz |  | 0.009997 | 700 |  |  |  | 20 |  | 33.7164
0.01 V, 55 Hz |  | 0.0099978 | 700 |  |  |  | 20 |  | 29.1612
0.01 V, 300 Hz |  | 0.0099999 | 700 |  |  |  | 20 |  | 58.7373
0.01 V, 1 KHz |  | 0.0099989 | 700 |  |  |  | 20 |  | 88.61
0.01 V, 10 KHz |  | 0.0099995 | 700 |  |  |  | 20 |  | 49.528
0.01 V, 20 KHz |  | 0.010001 | 700 |  |  |  | 20 |  | 43.8726
0.01 V, 30 KHz |  | 0.010002 | 940 |  |  |  | 20 |  | 22.9691
0.01 V, 50 KHz |  | 0.010006 | 940 |  |  |  | 20 |  | 40.595
0.01 V, 100 KHz |  | 0.010016 | 1600 |  |  |  | 50 |  | 25.831
0.01 V, 200 KHz |  | 0.010012 | 2600 |  |  |  | 100 |  | 25.025
0.01 V, 300 KHz |  | 0.010004 | 2600 |  |  |  | 100 |  | 20.6081
0.01 V, 500 KHz |  | 0.0099098 | 4500 |  |  |  | 200 |  | 21.4063
0.01 V, 1 MHz |  | 0.0086406 | 6000 |  |  |  | 300 |  | 28.6099
0.1 V, 10 Hz |  | 0.099731 | 660 |  |  |  | 20 |  | 16.7559
0.1 V, 20 Hz |  | 0.099928 | 300 |  |  |  | 20 |  | 2.6
0.1 V, 30 Hz |  | 0.099966 | 300 |  |  |  | 20 |  | 3.047
0.1 V, 40 Hz |  | 0.099979 | 195 |  |  |  | 20 |  | 4.7707
0.1 V, 55 Hz |  | 0.099987 | 195 |  |  |  | 20 |  | 6.8889
0.1 V, 300 Hz |  | 0.099995 | 195 |  |  |  | 20 |  | 8.6044
0.1 V, 1 KHz |  | 0.099996 | 195 |  |  |  | 20 |  | 4.5554
0.1 V, 10 KHz |  | 0.099991 | 195 |  |  |  | 20 |  | 5.4114
0.1 V, 20 KHz |  | 0.099986 | 195 |  |  |  | 20 |  | 5.6821
0.1 V, 30 KHz |  | 0.099985 | 400 |  |  |  | 20 |  | 5.601
0.1 V, 50 KHz |  | 0.099988 | 400 |  |  |  | 20 |  | 3.3462
0.1 V, 100 KHz |  | 0.10001 | 1050 |  |  |  | 50 |  | 3.082
0.1 V, 200 KHz |  | 0.099927 | 1240 |  |  |  | 100 |  | 3.3232
0.1 V, 300 KHz |  | 0.09967 | 1240 |  |  |  | 100 |  | 3.3444
0.1 V, 500 KHz |  | 0.098385 | 1900 |  |  |  | 200 |  | 1.6855
0.1 V, 1 MHz |  | 0.086887 | 4000 |  |  |  | 300 |  | 2.9963
1 V, 10 Hz |  | 0.99732 | 600 |  |  |  | 10 |  | 28.2714
1 V, 20 Hz |  | 0.99928 | 180 |  |  |  | 10 |  | 1.2702
1 V, 30 Hz |  | 0.99965 | 180 |  |  |  | 10 |  | 2.9591
1 V, 40 Hz |  | 0.99978 | 77 |  |  |  | 10 |  | 2.3516
1 V, 55 Hz |  | 0.99986 | 77 |  |  |  | 10 |  | 1.5002
1 V, 300 Hz |  | 0.99995 | 77 |  |  |  | 10 |  | 2.007
1 V, 1 KHz |  | 0.99996 | 77 |  |  |  | 10 |  | 1.7639
1 V, 10 KHz |  | 0.9999 | 77 |  |  |  | 10 |  | 0.441
1 V, 20 KHz |  | 0.99984 | 77 |  |  |  | 10 |  | 1.5368
1 V, 30 KHz |  | 0.99984 | 140 |  |  |  | 10 |  | 1.202
1 V, 50 KHz |  | 0.99988 | 140 |  |  |  | 20 |  | 1.3946
1 V, 100 KHz |  | 0.99989 | 310 |  |  |  | 30 |  | 3.0188
1 V, 200 KHz |  | 0.99976 | 550 |  |  |  | 70 |  | 0.92818
1 V, 300 KHz |  | 0.99997 | 550 |  |  |  | 70 |  | 1.59
1 V, 500 KHz |  | 1.0012 | 1400 |  |  |  | 100 |  | 1.4981
1 V, 1 MHz |  | 1.0042 | 3000 |  |  |  | 200 |  | 3.3442
10 V, 10 Hz |  | 9.9733 | 600 |  |  |  | 10 |  | 22.7604
10 V, 20 Hz |  | 9.9928 | 180 |  |  |  | 10 |  | 2.8647
10 V, 30 Hz |  | 9.9965 | 180 |  |  |  | 10 |  | 1.9657
10 V, 40 Hz |  | 9.9978 | 77 |  |  |  | 10 |  | 2.5061
10 V, 55 Hz |  | 9.9986 | 77 |  |  |  | 10 |  | 1.9652
10 V, 300 Hz |  | 9.9995 | 77 |  |  |  | 10 |  | 1.6668
10 V, 1 KHz |  | 9.9996 | 77 |  |  |  | 10 |  | 1.2019
10 V, 10 KHz |  | 9.9997 | 77 |  |  |  | 10 |  | 0.97186
10 V, 20 KHz |  | 9.9994 | 77 |  |  |  | 10 |  | 0.92802
10 V, 30 KHz |  | 9.9991 | 140 |  |  |  | 10 |  | 0.882
10 V, 50 KHz |  | 9.9984 | 140 |  |  |  | 20 |  | 1.5368
10 V, 100 KHz |  | 9.9966 | 270 |  |  |  | 30 |  | 1.3338
10 V, 200 KHz |  | 9.9921 | 670 |  |  |  | 70 |  | 1.0549
10 V, 300 KHz |  | 9.9893 | 670 |  |  |  | 70 |  | 0.97286
10 V, 500 KHz |  | 9.999 | 1700 |  |  |  | 100 |  | 1.7402
10 V, 1 MHz |  | 10.1282 | 3500 |  |  |  | 200 |  | 6.9075
19 V, 10 Hz |  | 18.953 | 552.63 |  |  |  | 10 |  | 19.3267
19 V, 20 Hz |  | 18.9874 | 165.79 |  |  |  | 10 |  | 1.2629
19 V, 30 Hz |  | 18.9938 | 165.79 |  |  |  | 10 |  | 2.4428
19 V, 40 Hz |  | 18.996 | 73.684 |  |  |  | 10 |  | 1.7744
19 V, 55 Hz |  | 18.9974 | 73.684 |  |  |  | 10 |  | 1.5145
19 V, 300 Hz |  | 18.9991 | 73.684 |  |  |  | 10 |  | 2.0215
19 V, 1 KHz |  | 18.9992 | 73.684 |  |  |  | 10 |  | 1.1996
19 V, 10 KHz |  | 18.9992 | 73.684 |  |  |  | 10 |  | 0.65057
19 V, 20 KHz |  | 18.9986 | 73.684 |  |  |  | 10 |  | 0.76477
19 V, 30 KHz |  | 18.9978 | 130.53 |  |  |  | 10 |  | 1.6248
19 V, 50 KHz |  | 18.9962 | 130.53 |  |  |  | 20 |  | 2.0632
19 V, 100 KHz |  | 18.9918 | 251.05 |  |  |  | 30 |  | 1.7856
19 V, 200 KHz |  | 18.9803 | 589.47 |  |  |  | 70 |  | 0.76551
19 V, 300 KHz |  | 18.9724 | 589.47 |  |  |  | 70 |  | 1.0245
19 V, 500 KHz |  | 18.9844 | 1463.2 |  |  |  | 100 |  | 1.3458
19 V, 1 MHz |  | 19.2042 | 3073.7 |  |  |  | 200 |  | 7.5224
100 V, 10 Hz |  | 99.7313 | 600 |  |  |  | 10 |  | 25.3306
100 V, 20 Hz |  | 99.9286 | 180 |  |  |  | 10 |  | 2.1294
100 V, 30 Hz |  | 99.9655 | 180 |  |  |  | 10 |  | 1.4245
100 V, 40 Hz |  | 99.978 | 85 |  |  |  | 10 |  | 1.3336
100 V, 55 Hz |  | 99.9859 | 85 |  |  |  | 10 |  | 1.0001
100 V, 300 Hz |  | 99.9951 | 85 |  |  |  | 10 |  | 1.4143
100 V, 1 KHz |  | 99.9956 | 85 |  |  |  | 10 |  | 1.2694
100 V, 10 KHz |  | 99.9889 | 85 |  |  |  | 10 |  | 1.202
100 V, 20 KHz |  | 99.9877 | 85 |  |  |  | 10 |  | 2.2916
100 V, 30 KHz |  | 99.9853 | 240 |  |  |  | 10 |  | 1.202
100 V, 50 KHz |  | 99.9781 | 240 |  |  |  | 20 |  | 1.4533
100 V, 100 KHz |  | 99.9458 | 600 |  |  |  | 30 |  | 5.1856
100 V, 200 KHz |  | 99.8546 | 2600 |  |  |  | 50 |  | 13.5516
1000 V, 55 Hz |  | 1009.721 | 79 |  |  |  | 15 |  | 1.9669
1000 V, 300 Hz |  | 1009.83 | 79 |  |  |  | 15 |  | 1.1194
700 V, 1 KHz |  | 709.866 | 80.714 |  |  |  | 50 |  | 1.7253
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.9915e-07 | 10053.5455 | [5 empty] | 57.7841
10 µADC | 1e-05 | 9.999e-06 | 1045.1097 | [5 empty] | 5.7741
50 µADC | 5e-05 | 4.9998e-05 | 245 | [5 empty] | 1
100 µADC | 0.0001 | 9.9997e-05 | 145 |  |  |  | 7 |  | 0.50001
-100 µADC | -0.0001 | -9.9999e-05 | 145 |  |  |  | 7 |  | -0.5
-50 µADC | -5e-05 | -5e-05 | 245 | [5 empty] | -2
-10 µADC | -1e-05 | -1.0001e-05 | -1044.9425 | [5 empty] | -5.7732
0.1 mADC | 0.0001 | 0.00010008 | 145 | [5 empty] | 5.7689
0.5 mADC | 0.0005 | 0.00050004 | 65 | [5 empty] | 0.99992
1 mADC | 0.001 | 0.00099999 | 55 |  |  |  | 7 |  | 0.50001
-1 mADC | -0.001 | -0.00099999 | 55 |  |  |  | 7 |  | -0.50001
-0.5 mADC | -0.0005 | -0.00049995 | 65 | [5 empty] | -1.0001
1 mADC | 0.001 | 0.0010008 | 55 | [5 empty] | 4.9961
5 mADC | 0.005 | 0.0050004 | 65 | [5 empty] | 0.99992
10 mADC | 0.01 | 0.0099999 | 55 |  |  |  | 7 |  | 0.50001
-10 mADC | -0.01 | -0.0099999 | 55 |  |  |  | 7 |  | -0.50001
-5 mADC | -0.005 | -0.0049995 | 65 | [5 empty] | 0
10 mADC | 0.01 | 0.010008 | 55 | [5 empty] | 0
50 mADC | 0.05 | 0.050004 | 75 | [5 empty] | 1.9998
100 mADC | 0.1 | 0.099998 | 65 |  |  |  | 7 |  | 0
-100 mADC | -0.1 | -0.099999 | 65 |  |  |  | 7 |  | -0.50001
-50 mADC | -0.05 | -0.049995 | 75 | [5 empty] | -1.0001
0.1 ADC | 0.1 | 0.10009 | 65 | [5 empty] | 4.9955
0.5 ADC | 0.5 | 0.50004 | 135 | [5 empty] | 0
-0.5 ADC | -0.5 | -0.49995 | 135 | [5 empty] | -1.0001
-1.0 ADC | -1 | -1 | 105 |  |  |  | 20 |  | -0.57735
1.0 ADC | 1 | 0.99996 | 105 |  |  |  | 20 |  | 0.50002
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.978e-06 | 3658.0477 | [5 empty] | 50.1102
10 µA, 20 Hz | 1e-05 | 9.997e-06 | 2850.8553 | [5 empty] | 0
10 µA, 30 Hz | 1e-05 | 1.0001e-05 | 2849.715 | [5 empty] | 0
10 µA, 40 Hz | 1e-05 | 1.0002e-05 | 2119.5761 | [5 empty] | 49.99
10 µA 55 Hz | 1e-05 | 1.0002e-05 | 2119.4701 | [5 empty] | 57.7206
10 µA, 300 Hz | 1e-05 | 1.0004e-05 | 2119.1523 |  |  | -40 |  |  | 125.7803
10 µA, 1 KHz | 1e-05 | 1.0004e-05 | 2119.1523 | [5 empty] | 99.96
10 µA, 5 KHz | 1e-05 | 9.9995e-06 | 5500.275 | [5 empty] | 95.7475
10 µA, 10 KHz | 1e-05 | 9.981e-06 | 11521.8916 | [5 empty] | 100.1904
100 µA, 10 Hz | 0.0001 | 9.9757e-05 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 38.8244
100 µA, 20 Hz | 0.0001 | 9.9943e-05 | 600 |  |  |  | 50 |  | 5.7768
100 µA, 30 Hz | 0.0001 | 9.9979e-05 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 5.0011
100 µA, 40 Hz | 0.0001 | 9.9991e-05 | 320 |  |  |  | 50 |  | 5.774
100 µA 55 Hz | 0.0001 | 9.9999e-05 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 5.0001
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 17.3196
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 9.574
100 µA, 5 KHz | 0.0001 | 9.996e-05 | 0.1 |  |  |  | 100 |  | 5.002
100 µA, 10 KHz | 0.0001 | 9.9821e-05 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 8.1796
1mA, 10 Hz | 0.001 | 0.00099774 | 700 |  |  |  | 40 |  | 43.5916
1 mA, 20 Hz | 0.001 | 0.00099961 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 5.002
1 mA, 30 Hz | 0.001 | 0.00099997 | 390 |  |  |  | 40 |  | 5.0002
1 mA, 40 Hz | 0.001 | 0.0010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0
1 mA, 55 Hz | 0.001 | 0.0010002 | 160 |  |  |  | 40 |  | 4.9992
1 mA, 300 Hz | 0.001 | 0.0010003 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 9.5717
1 mA, 1 KHz | 0.001 | 0.0010002 | 160 |  |  |  | 40 |  | 9.5721
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 5.7733
1 mA, 10 KHz | 0.001 | 0.00099943 | 0.25 |  |  |  | 200 |  | 8.1696
10 mA, 10 Hz | 0.01 | 0.0099776 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 15.0337
10 mA, 20 Hz | 0.01 | 0.009996 | 390 |  |  |  | 40 |  | 5.7758
10 mA, 30 Hz | 0.01 | 0.0099997 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 5.0002
10 mA, 40 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 4.9996
10 mA, 55 Hz | 0.01 | 0.010002 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 4.9992
10 mA, 300 Hz | 0.01 | 0.010002 | 160 |  |  |  | 40 |  | 9.5719
10 mA, 1 KHz | 0.01 | 0.010002 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 5.7722
10 mA, 5 KHz | 0.01 | 0.01 | 1000 |  |  |  | 70 |  | 12.9095
10 mA, 10 KHz | 0.01 | 0.0099948 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 5.0026
100 mA, 10 Hz | 0.1 | 0.099766 | 700 |  |  |  | 40 |  | 12.9403
100 mA, 20 Hz | 0.1 | 0.09995 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0
100 mA, 30 Hz | 0.1 | 0.099986 | 390 |  |  |  | 40 |  | 0
100 mA, 40 Hz | 0.1 | 0.1 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 5
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 4.9997
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 12.5813
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 9.5734
100 mA, 5 KHz | 0.1 | 0.1 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 9.5739
100 mA, 10 KHz | 0.1 | 0.099971 | 0.25 |  |  |  | 200 |  | 5.7752
1 A, 10 Hz | 1 | 0.99828 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 16.3581
1 A, 20 Hz | 1 | 1.0002 | -1 |  |  |  | 40 |  | 0
1 A, 30 Hz | 1 | 1.0005 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0
1 A, 40 Hz | 1 | 1.0006 | 640 |  |  |  | 40 |  | 0
1 A, 55 Hz | 1 | 1.0007 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 8.159
1 A, 300 Hz | 1 | 1.0008 | 640 |  |  |  | 40 |  | 14.9882
1 A, 1 KHz | 1 | 1.0005 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 9.5693
1 A, 5 KHz | 1 | 1.0001 | 800 |  |  |  | 70 |  | 20.613
1 A, 10 KHz | 1 | 0.99863 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 9.5874
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.01 | 277.45 | 280.04 | 66.65 | 20.28 | 23.3