## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 43268
Model Number |  | 1281 |  |  | Test date |  | 24-10-15 21:30:57
Meter Info |  |  |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1,1,0,0,0 |  |  | CERT_AMB?
Firmware |  | 890144-03.14 |  |  | TEMP?
Date of Baseline Cal |  |  |  |  | TEMP? BASE
Date of Certified Cal |  |  |  |  | TEMP? CERTIFIED
CNFTST? |  | 0 |  |  | STANDBY?
DUT Confidence level
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 290524 |  |  | MFC last calibrated |  | 139 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 290524 |  |  | MFC since DCV Zero |  | 139 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529890647135340E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306130764516422E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000045271113747E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000069664541200E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996126071898848E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6066 days 5.67 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 3.55 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.042 | 7.5e-07 |  |  | N/A |  |  | 0.046296 | 0.035
0.0VDC | 0 | -0.52 | 7.5e-07 |  |  | N/A |  |  | 0.056593 | -0.48
0.0VDC | 0 | -15.1 | 7.5e-07 |  |  | N/A |  |  | 0.24495 | -15.4
0.0VDC | 0 | -127 | 7.5e-07 |  |  | N/A |  |  | 4.5856 | -124
0.0VDC | 0 | -260 | 7.5e-07 |  |  | N/A |  |  | 39.0512 | -190
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.0099999 | 81.5 | [5 empty] | 27.4107 | 0.0099999
0.02 | 100mV | 0.02 | 44 | [5 empty] | 2.0646 | 0.02
0.05 | 100mV | 0.049999 | 21.5 | [5 empty] | 4.2931 | 0.049999
0.1 | 100mV | 0.099999 | 14 |  |  |  | 3 |  | 0.71607 | 0.099999
-0.01 | 100mV | -0.0099999 | -81.5 | [5 empty] | -7.12 | -0.0099997
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -4.2035 | -0.02
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -1.37 | -0.05
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.85193 | -0.1
0.1 | 1V | 0.099999 | 14 | [5 empty] | 1.006 | 0.099999
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.25386 | 0.2
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.37821 | 0.5
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.53175 | 1
-0.1 | 1V | -0.1 | 14 | [5 empty] | -0.47463 | -0.1
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.45697 | -0.2
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.097525 | -0.5
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.094531 | -1
1 | 10V | 0.99998 | 4.7 | [5 empty] | 0.52679 | 0.99998
2 | 10V | 2 | 4.1 | [5 empty] | 0.19526 | 2
5 | 10V | 5 | 4.1 | [5 empty] | 0.069841 | 5
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.025712 | 10
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.029067 | 19
-1 | 10V | -1 | 4.7 | [5 empty] | -0.80466 | -1
-2 | 10V | -2 | 4.1 | [5 empty] | -0.21618 | -2
-5 | 10V | -5 | 4.1 | [5 empty] | -0.12369 | -5
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.039299 | -10
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.034989 | -19
10 | 100V | 9.9998 | 3.8 | [5 empty] | 0.73656 | 9.9998
20 | 100V | 19.9998 | 3.8 | [5 empty] | 0.18503 | 19.9998
50 | 100V | 49.9997 | 7 | [5 empty] | 0.13606 | 49.9997
100 | 100V | 99.9995 | 6 |  |  |  | 2 |  | 0.041966 | 99.9995
-10 | 100V | -10.0001 | 3.8 | [5 empty] | -0.48419 | -10.0001
-20 | 100V | -20.0001 | 3.8 | [5 empty] | -0.20156 | -20.0001
-50 | 100V | -50 | 7 | [5 empty] | -0.11681 | -50
-100 | 100V | -99.9998 | 6 |  |  |  | 2 |  | -0.054722 | -99.9998
-100 | 1000V | -99.9998 | 6 | [5 empty] | -0.62472 | -99.9998
-200 | 1000V | -199.9993 | 5.5 | [5 empty] | -0.1965 | -199.9993
-500 | 1000V | -499.9978 | 8.2 | [5 empty] | -0.099889 | -499.9978
-1000 | 1000V | -999.9938 | 7.6 |  |  |  | 2 |  | -0.039406 | -999.9938
100 | 1000V | 99.999 | 6 | [5 empty] | 0.1716 | 99.999
200 | 1000V | 199.9983 | 5.5 | [5 empty] | 0.12937 | 199.9983
500 | 1000V | 499.9965 | 8.2 | [5 empty] | 0.072189 | 499.9965
1000 | 1000V | 999.9921 | 7.6 |  |  |  | 2 |  | 0.046756 | 999.992
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99961 | 0.9997 | 85 |  |  |  | 20 |  | 1.1779e-10 | 0.9997
1.9 Ω | 1.9 | 1.9 | 85 |  |  |  | 15 |  | 17.5439 | 1.9
10 Ω | 10.0004 | 10.0004 | 26 |  |  |  | 5 |  | 0 | 10.0004
19 Ω | 19.0039 | 19.0039 | 24 |  |  |  | 5 |  | 0 | 19.0039
100 Ω | 100.0053 | 100.0048 | 15 |  |  |  | 3 |  | 0 | 100.0048
190 Ω | 190.0032 | 190.0026 | 15 |  |  |  | 3 |  | 0 | 190.0026
1 kΩ | 999.9469 | 999.932 | 11 |  |  |  | 3 |  | 0.44099 | 999.932
1.9 kΩ | 1899.9106 | 1899.887 | 11 |  |  |  | 3 |  | 0.2321 | 1899.887
10 kΩ | 10000.453 | 10000.36 | 9 |  |  |  | 3 |  | 0.52703 | 10000.36
19 kΩ | 18999.722 | 18999.57 | 9 |  |  |  | 3 |  | 0.37217 | 18999.57
100 kΩ | 99998.75 | 99998.2 | 11 |  |  |  | 5 |  | 0.50001 | 99998.2
190 kΩ | 189994.66 | 189993.7 | 11 |  |  |  | 5 |  | 0.7895 | 189993.6
1 MΩ | 999908.7 | 999895 | 16 |  |  |  | 8 |  | 2.9062 | 999895
1.9 MΩ | 1899855.1 | 1899828 | 17 |  |  |  | 8 |  | 0.64466 | 1899829
10MΩ | 9998564 | 9999120 | 33 |  |  |  | 12 |  | 1.6416 | 9999100
19 MΩ | 18999500 | 18999890 | 43 |  |  |  | 12 |  | 1.3274 | 18999900
100MΩ(2W) | 100009390 | 2e+35 | 110 |  |  |  | 180 |  | 0 | 2e+35
| 100009390 | 2e+35 | 110 | [5 empty] | 0 | 2e+35
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0 | 5e-05 |  |  | N/A |  |  | 0 | 0
0 Ω | 100 Ω | 0 | 5e-05 |  |  | N/A |  |  | 3.3333e-05 | 0
0 Ω | 1 KΩ | 0 | 5e-05 |  |  | N/A |  |  | 3.3333e-05 | 0
0 Ω | 10 KΩ | -0.002 | 5e-05 |  |  | N/A |  |  | 0.00033333 | -0.002
0 Ω | 100 KΩ | -0.02 | 5e-05 |  |  | N/A |  |  | 0.0044096 | -0.01
0 Ω | 1 MΩ | -0.1 | 5e-05 |  |  | N/A |  |  | 0 | -0.1
0 Ω | 10 MΩ | -1 | 5e-05 |  |  | N/A |  |  | 0 | -1
0 Ω | 100 MΩ(2W) | -10 | 5e-05 |  |  | N/A |  |  | 3.3333 | -10
|  | -10 | 5e-05 | [5 empty] | 0 | -10
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 1.2365 | 5e-05 |  |  | N/A |  |  | 3.3333e-05 | 1.2365
0 Ω | 100 Ω | 1.2365 | 5e-05 |  |  | N/A |  |  | 0 | 1.2365
0 Ω | 1 KΩ | 1.2365 | 5e-05 |  |  | N/A |  |  | 0 | 1.2365
0 Ω | 10 KΩ | 1.235 | 5e-05 |  |  | N/A |  |  | 0 | 1.235
0 Ω | 100 KΩ | 1.22 | 5e-05 |  |  | N/A |  |  | 0 | 1.22
0 Ω | 1 MΩ | 1.2 | 5e-05 |  |  | N/A |  |  | 0 | 1.2
0 Ω | 10 MΩ | 1 | 5e-05 |  |  | N/A |  |  | 0.44096 | 1
0 Ω | 100 MΩ | -10 | 5e-05 |  |  | N/A |  |  | 0 | -10
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99869 | 77 |  |  |  | 10 |  | 25.5285 | 0.9987
2VAC/1kHz | 10 VAC | 1.9973 | 73.5 |  |  |  | 10 |  | 14.8337 | 1.9973
5VAC/1kHz | 10 VAC | 4.9932 | 84 |  |  |  | 10 |  | 4.1824 | 4.9931
10VAC/1kHz | 10 VAC | 9.9862 | 77 |  |  |  | 10 |  | 2.2267 | 9.9862
12VAC/1kHz | 10 VAC | 11.9834 | 75.833 |  |  |  | 10 |  | 2.2078 | 11.9834
15VAC/1kHz | 10 VAC | 14.9791 | 74.667 |  |  |  | 10 |  | 2.2336 | 14.9791
19VAC/1kHz | 10 VAC | 18.9735 | 73.684 |  |  |  | 10 |  | 1.1915 | 18.9734
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0023004 | 5500 |  |  |  | 20 |  | 7142.4359 | 0.002295
0.001 V, 20 Hz |  | 0.0023165 | 5200 |  |  |  | 20 |  | 6502.9786 | 0.0023176
0.001 V, 30 Hz |  | 0.0023668 | 5200 |  |  |  | 20 |  | 20206.7489 | 0.002371
0.001 V, 40 Hz |  | 0.0023577 | 5100 |  |  |  | 20 |  | 62309.1454 | 0.0024477
0.001 V, 55 Hz |  | 0.0024141 | 5100 |  |  |  | 20 |  | 163516.9844 | 0.0022022
0.001 V, 300 Hz |  | 0.0022948 | 5100 |  |  |  | 20 |  | 4522.2329 | 0.0022953
0.001 V, 1 KHz |  | 0.0023064 | 5100 |  |  |  | 20 |  | 3147.5358 | 0.0022838
0.001 V, 10 KHz |  | 0.0022761 | 5100 |  |  |  | 20 |  | 1584.7009 | 0.0022807
0.001 V, 20 KHz |  | 0.0022773 | 5100 |  |  |  | 20 |  | 2231.5636 | 0.0022774
0.001 V, 30 KHz |  | 0.0022726 | 5340 |  |  |  | 20 |  | 1397.9009 | 0.0022702
0.001 V, 50 KHz |  | 0.0022846 | 5340 |  |  |  | 20 |  | 1545.9161 | 0.0022797
0.001 V, 100 KHz |  | 0.0022783 | 8800 |  |  |  | 50 |  | 1303.5371 | 0.0022896
0.001 V, 200 KHz |  | 0.0022713 | 16100 |  |  |  | 100 |  | 1561.9132 | 0.0022711
0.001 V, 300 KHz |  | 0.0022606 | 16100 |  |  |  | 100 |  | 1730.1767 | 0.0022576
0.001 V, 500 KHz |  | 0.0022502 | 31500 |  |  |  | 200 |  | 2094.6045 | 0.0022533
0.001 V, 1 MHz |  | 0.0021786 | 33000 |  |  |  | 300 |  | 2120.7129 | 0.0021764
0.01 V, 10 Hz |  | 0.010785 | 1100 |  |  |  | 20 |  | 28758.3909 | 0.010866
0.01 V, 20 Hz |  | 0.010133 | 800 |  |  |  | 20 |  | 3103.0334 | 0.01016
0.01 V, 30 Hz |  | 0.010175 | 800 |  |  |  | 20 |  | 11777.5349 | 0.010089
0.01 V, 40 Hz |  | 0.010309 | 700 |  |  |  | 20 |  | 31143.1952 | 0.010249
0.01 V, 55 Hz |  | 0.010323 | 700 |  |  |  | 20 |  | 98467.7906 | 0.0088049
0.01 V, 300 Hz |  | 0.010158 | 700 |  |  |  | 20 |  | 4089.605 | 0.010147
0.01 V, 1 KHz |  | 0.010122 | 700 |  |  |  | 20 |  | 1980.3173 | 0.010147
0.01 V, 10 KHz |  | 0.0094926 | 700 |  |  |  | 20 |  | 250.1317 | 0.0094897
0.01 V, 20 KHz |  | 0.0093951 | 700 |  |  |  | 20 |  | 461.3664 | 0.0093825
0.01 V, 30 KHz |  | 0.0093693 | 940 |  |  |  | 20 |  | 150.6054 | 0.0093697
0.01 V, 50 KHz |  | 0.0093938 | 940 |  |  |  | 20 |  | 354.8706 | 0.0093992
0.01 V, 100 KHz |  | 0.0094357 | 1600 |  |  |  | 50 |  | 73.4467 | 0.0094369
0.01 V, 200 KHz |  | 0.0094878 | 2600 |  |  |  | 100 |  | 118.479 | 0.0094888
0.01 V, 300 KHz |  | 0.0095761 | 2600 |  |  |  | 100 |  | 72.1812 | 0.0095755
0.01 V, 500 KHz |  | 0.009773 | 4500 |  |  |  | 200 |  | 214.6814 | 0.0097701
0.01 V, 1 MHz |  | 0.0089586 | 6000 |  |  |  | 300 |  | 129.696 | 0.0089593
0.1 V, 10 Hz |  | 0.10794 | 660 |  |  |  | 20 |  | 33701.208 | 0.10962
0.1 V, 20 Hz |  | 0.099891 | 300 |  |  |  | 20 |  | 1543.6628 | 0.10009
0.1 V, 30 Hz |  | 0.1002 | 300 |  |  |  | 20 |  | 1263.1614 | 0.10028
0.1 V, 40 Hz |  | 0.10021 | 195 |  |  |  | 20 |  | 3327.2923 | 0.1001
0.1 V, 55 Hz |  | 0.099856 | 195 |  |  |  | 20 |  | 10173.5833 | 0.10164
0.1 V, 300 Hz |  | 0.10023 | 195 |  |  |  | 20 |  | 220.3705 | 0.10028
0.1 V, 1 KHz |  | 0.099913 | 195 |  |  |  | 20 |  | 219.8574 | 0.099955
0.1 V, 10 KHz |  | 0.093394 | 195 |  |  |  | 20 |  | 29.1467 | 0.093391
0.1 V, 20 KHz |  | 0.092329 | 195 |  |  |  | 20 |  | 46.4718 | 0.092334
0.1 V, 30 KHz |  | 0.092096 | 400 |  |  |  | 20 |  | 17.8786 | 0.092097
0.1 V, 50 KHz |  | 0.092292 | 400 |  |  |  | 20 |  | 17.5531 | 0.092294
0.1 V, 100 KHz |  | 0.092606 | 1050 |  |  |  | 50 |  | 5.9826 | 0.092606
0.1 V, 200 KHz |  | 0.093174 | 1240 |  |  |  | 100 |  | 16.6769 | 0.093173
0.1 V, 300 KHz |  | 0.093898 | 1240 |  |  |  | 100 |  | 4.4232 | 0.093898
0.1 V, 500 KHz |  | 0.094934 | 1900 |  |  |  | 200 |  | 11.8955 | 0.094933
0.1 V, 1 MHz |  | 0.084795 | 4000 |  |  |  | 300 |  | 16.8932 | 0.084796
1 V, 10 Hz |  | 1.0947 | 600 |  |  |  | 10 |  | 27176.0847 | 1.0739
1 V, 20 Hz |  | 0.99707 | 180 |  |  |  | 10 |  | 1596.6726 | 0.99663
1 V, 30 Hz |  | 1.0004 | 180 |  |  |  | 10 |  | 281.5469 | 1.0006
1 V, 40 Hz |  | 1.0018 | 77 |  |  |  | 10 |  | 278.8489 | 1.0017
1 V, 55 Hz |  | 1.0016 | 77 |  |  |  | 10 |  | 1016.8433 | 1.0033
1 V, 300 Hz |  | 1.002 | 77 |  |  |  | 10 |  | 49.4153 | 1.002
1 V, 1 KHz |  | 0.99861 | 77 |  |  |  | 10 |  | 29.4067 | 0.99856
1 V, 10 KHz |  | 0.93354 | 77 |  |  |  | 10 |  | 4.5901 | 0.93354
1 V, 20 KHz |  | 0.92293 | 77 |  |  |  | 10 |  | 6.1875 | 0.92292
1 V, 30 KHz |  | 0.92074 | 140 |  |  |  | 10 |  | 1.9412 | 0.92074
1 V, 50 KHz |  | 0.92285 | 140 |  |  |  | 20 |  | 2.4828 | 0.92285
1 V, 100 KHz |  | 0.9259 | 310 |  |  |  | 30 |  | 4.1868 | 0.92591
1 V, 200 KHz |  | 0.93197 | 550 |  |  |  | 70 |  | 1.6194 | 0.93198
1 V, 300 KHz |  | 0.94011 | 550 |  |  |  | 70 |  | 1.1625 | 0.94011
1 V, 500 KHz |  | 0.96196 | 1400 |  |  |  | 100 |  | 3.075 | 0.96196
1 V, 1 MHz |  | 0.9791 | 3000 |  |  |  | 200 |  | 2.0709 | 0.9791
10 V, 10 Hz |  | 10.6421 | 600 |  |  |  | 10 |  | 34790.6677 | 10.3265
10 V, 20 Hz |  | 9.9871 | 180 |  |  |  | 10 |  | 1622.3301 | 9.9629
10 V, 30 Hz |  | 10.0031 | 180 |  |  |  | 10 |  | 230.1903 | 10.002
10 V, 40 Hz |  | 10.0191 | 77 |  |  |  | 10 |  | 71.1178 | 10.0191
10 V, 55 Hz |  | 10.023 | 77 |  |  |  | 10 |  | 101.5493 | 10.0219
10 V, 300 Hz |  | 10.0201 | 77 |  |  |  | 10 |  | 5.7331 | 10.02
10 V, 1 KHz |  | 9.9861 | 77 |  |  |  | 10 |  | 1.5921 | 9.9861
10 V, 10 KHz |  | 9.3351 | 77 |  |  |  | 10 |  | 1.4937 | 9.3351
10 V, 20 KHz |  | 9.2276 | 77 |  |  |  | 10 |  | 2.5415 | 9.2276
10 V, 30 KHz |  | 9.2085 | 140 |  |  |  | 10 |  | 1.8896 | 9.2085
10 V, 50 KHz |  | 9.2314 | 140 |  |  |  | 20 |  | 3.2246 | 9.2314
10 V, 100 KHz |  | 9.2614 | 270 |  |  |  | 30 |  | 1.3224 | 9.2614
10 V, 200 KHz |  | 9.3276 | 670 |  |  |  | 70 |  | 2.0373 | 9.3276
10 V, 300 KHz |  | 9.4142 | 670 |  |  |  | 70 |  | 1.3827 | 9.4142
10 V, 500 KHz |  | 9.6479 | 1700 |  |  |  | 100 |  | 1.476 | 9.6479
10 V, 1 MHz |  | 9.9283 | 3500 |  |  |  | 200 |  | 7.6505 | 9.9282
19 V, 10 Hz |  | 2e+35 | 552.63 |  |  |  | 10 |  | 333333.3333 | 2e+35
19 V, 20 Hz |  | 18.9607 | 165.79 |  |  |  | 10 |  | 1681.9203 | 18.9982
19 V, 30 Hz |  | 19.0049 | 165.79 |  |  |  | 10 |  | 214.1495 | 19.0026
19 V, 40 Hz |  | 19.037 | 73.684 |  |  |  | 10 |  | 74.9327 | 19.0339
19 V, 55 Hz |  | 19.0433 | 73.684 |  |  |  | 10 |  | 54.3817 | 19.0422
19 V, 300 Hz |  | 19.0379 | 73.684 |  |  |  | 10 |  | 2.4419 | 19.0378
19 V, 1 KHz |  | 18.9733 | 73.684 |  |  |  | 10 |  | 1.7101 | 18.9732
19 V, 10 KHz |  | 17.7352 | 73.684 |  |  |  | 10 |  | 1.3811 | 17.7352
19 V, 20 KHz |  | 17.521 | 73.684 |  |  |  | 10 |  | 1.2032 | 17.5211
19 V, 30 KHz |  | 17.5148 | 130.53 |  |  |  | 10 |  | 2.4465 | 17.5148
19 V, 50 KHz |  | 17.544 | 130.53 |  |  |  | 20 |  | 0.89117 | 17.544
19 V, 100 KHz |  | 17.594 | 251.05 |  |  |  | 30 |  | 0.73986 | 17.594
19 V, 200 KHz |  | 17.7156 | 589.47 |  |  |  | 70 |  | 1.2302 | 17.7156
19 V, 300 KHz |  | 17.8753 | 589.47 |  |  |  | 70 |  | 1.0466 | 17.8753
19 V, 500 KHz |  | 18.3021 | 1463.2 |  |  |  | 100 |  | 1.3322 | 18.3021
19 V, 1 MHz |  | 18.7736 | 3073.7 |  |  |  | 200 |  | 2.3218 | 18.7735
100 V, 10 Hz |  | 106.7372 | 600 |  |  |  | 10 |  | 31140.2473 | 106.8886
100 V, 20 Hz |  | 99.7959 | 180 |  |  |  | 10 |  | 1713.3105 | 99.993
100 V, 30 Hz |  | 100.0293 | 180 |  |  |  | 10 |  | 222.2427 | 100.0398
100 V, 40 Hz |  | 100.1805 | 85 |  |  |  | 10 |  | 68.3748 | 100.1923
100 V, 55 Hz |  | 100.2256 | 85 |  |  |  | 10 |  | 10.9083 | 100.2233
100 V, 300 Hz |  | 100.1982 | 85 |  |  |  | 10 |  | 1.8965 | 100.198
100 V, 1 KHz |  | 99.8648 | 85 |  |  |  | 10 |  | 1.2264 | 99.865
100 V, 10 KHz |  | 93.3477 | 85 |  |  |  | 10 |  | 1.0713 | 93.3477
100 V, 20 KHz |  | 92.5419 | 85 |  |  |  | 10 |  | 1.3234 | 92.5419
100 V, 30 KHz |  | 92.3924 | 240 |  |  |  | 10 |  | 0.7863 | 92.3924
100 V, 50 KHz |  | 92.3738 | 240 |  |  |  | 20 |  | 2.0572 | 92.3738
100 V, 100 KHz |  | 92.5979 | 600 |  |  |  | 30 |  | 5.3756 | 92.5974
100 V, 200 KHz |  | 93.2595 | 2600 |  |  |  | 50 |  | 14.3605 | 93.2589
1000 V, 55 Hz |  | 974.391 | 79 |  |  |  | 15 |  | 5.6781 | 974.382
1000 V, 300 Hz |  | 974.117 | 79 |  |  |  | 15 |  | 2.3522 | 974.114
700 V, 1 KHz |  | 679.615 | 80.714 |  |  |  | 50 |  | 10.2999 | 679.602
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.9915e-07 | 10053.5455 | [5 empty] | 57.7841 | 9.99e-07
10 µADC | 1e-05 | 9.9993e-06 | 1045.0732 | [5 empty] | 0 | 9.9993e-06
50 µADC | 5e-05 | 5e-05 | 245 | [5 empty] | 0 | 5e-05
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0.81649 | 0.0001
-100 µADC | -0.0001 | -0.0001 | 145 |  |  |  | 7 |  | -0.57734 | -0.0001
-50 µADC | -5e-05 | -5.0001e-05 | 245 | [5 empty] | -0.99997 | -5.0001e-05
-10 µADC | -1e-05 | -1.0001e-05 | -1044.8903 | [5 empty] | -5.7729 | -1.0001e-05
0.1 mADC | 0.0001 | 9.9997e-05 | 145 | [5 empty] | 5.0002 | 9.9997e-05
0.5 mADC | 0.0005 | 0.00049999 | 65 | [5 empty] | 1.1547 | 0.00049999
1 mADC | 0.001 | 0.00099999 | 55 |  |  |  | 7 |  | 0.50001 | 0.00099999
-1 mADC | -0.001 | -0.00099999 | 55 |  |  |  | 7 |  | -0.57736 | -0.00099999
-0.5 mADC | -0.0005 | -0.0005 | 65 | [5 empty] | 0 | -0.0005
1 mADC | 0.001 | 0.0010001 | 55 | [5 empty] | 0 | 0.0010001
5 mADC | 0.005 | 0.0050004 | 65 | [5 empty] | 0.99992 | 0.0050004
10 mADC | 0.01 | 0.010001 | 55 |  |  |  | 7 |  | 0.95735 | 0.010001
-10 mADC | -0.01 | -0.010001 | 55 |  |  |  | 7 |  | -0.49996 | -0.010001
-5 mADC | -0.005 | -0.0050004 | 65 | [5 empty] | 0 | -0.0050004
10 mADC | 0.01 | 0.01 | 55 | [5 empty] | 4.9999 | 0.01
50 mADC | 0.05 | 0.050002 | 75 | [5 empty] | 0 | 0.050003
100 mADC | 0.1 | 0.10001 | 65 |  |  |  | 7 |  | 0.57732 | 0.10001
-100 mADC | -0.1 | -0.10001 | 65 |  |  |  | 7 |  | -0.49997 | -0.10001
-50 mADC | -0.05 | -0.050003 | 75 | [5 empty] | -1.1546 | -0.050003
0.1 ADC | 0.1 | 0.1 | 65 | [5 empty] | 5.7734 | 0.1
0.5 ADC | 0.5 | 0.5 | 135 | [5 empty] | 0.99999 | 0.5
-0.5 ADC | -0.5 | -0.50002 | 135 | [5 empty] | 0 | -0.50002
-1.0 ADC | -1 | -0.99996 | 105 |  |  |  | 20 |  | -0.57737 | -0.99996
1.0 ADC | 1 | 0.99992 | 105 |  |  |  | 20 |  | 0.81656 | 0.99992
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 1.1305e-05 | 3228.6599 | [5 empty] | 123585.794 | 1.2537e-05
10 µA, 20 Hz | 1e-05 | 9.93e-06 | 2870.0906 | [5 empty] | 6671.1477 | 9.95e-06
10 µA, 30 Hz | 1e-05 | 9.9865e-06 | 2853.8527 | [5 empty] | 1081.5836 | 9.986e-06
10 µA, 40 Hz | 1e-05 | 1.0008e-05 | 2118.3054 | [5 empty] | 3457.8454 | 9.982e-06
10 µA 55 Hz | 1e-05 | 9.99e-06 | 2122.1221 | [5 empty] | 3181.6246 | 9.985e-06
10 µA, 300 Hz | 1e-05 | 1.0006e-05 | 2118.6229 |  |  | -40 |  |  | 170.6716 | 1.0005e-05
10 µA, 1 KHz | 1e-05 | 1.0001e-05 | 2119.788 | [5 empty] | 341.5309 | 1.0001e-05
10 µA, 5 KHz | 1e-05 | 9.998e-06 | 5501.1002 | [5 empty] | 50.01 | 9.998e-06
10 µA, 10 KHz | 1e-05 | 9.981e-06 | 11521.8916 | [5 empty] | 173.5348 | 9.981e-06
100 µA, 10 Hz | 0.0001 | 0.00010117 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 155380.6761 | 0.00010478
100 µA, 20 Hz | 0.0001 | 9.9854e-05 | 600 |  |  |  | 50 |  | 7513.3506 | 9.9187e-05
100 µA, 30 Hz | 0.0001 | 9.9731e-05 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 482.5787 | 9.9903e-05
100 µA, 40 Hz | 0.0001 | 9.991e-05 | 320 |  |  |  | 50 |  | 341.5324 | 9.9909e-05
100 µA 55 Hz | 0.0001 | 9.9985e-05 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 266.0443 | 9.9943e-05
100 µA, 300 Hz | 0.0001 | 9.9989e-05 | 320 |  |  |  | 50 |  | 5.7741 | 9.9989e-05
100 µA, 1 KHz | 0.0001 | 9.9988e-05 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 25.1691 | 9.9985e-05
100 µA, 5 KHz | 0.0001 | 9.9946e-05 | 0.1 |  |  |  | 100 |  | 8.1694 | 9.9947e-05
100 µA, 10 KHz | 0.0001 | 9.9802e-05 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 10.0198 | 9.9802e-05
1mA, 10 Hz | 0.001 | 0.0011254 | 700 |  |  |  | 40 |  | 139481.1545 | 0.0011315
1 mA, 20 Hz | 0.001 | 0.00099318 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 7112.2536 | 0.0010033
1 mA, 30 Hz | 0.001 | 0.00099913 | 390 |  |  |  | 40 |  | 234.2099 | 0.00099841
1 mA, 40 Hz | 0.001 | 0.00099933 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 213.059 | 0.00099905
1 mA, 55 Hz | 0.001 | 0.00099992 | 160 |  |  |  | 40 |  | 33.0432 | 0.00099983
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 25.8199 | 0.00099999
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 18.2572 | 0.001
1 mA, 5 KHz | 0.001 | 0.00099997 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 9.5745 | 0.00099997
1 mA, 10 KHz | 0.001 | 0.00099967 | 0.25 |  |  |  | 200 |  | 9.5774 | 0.00099969
10 mA, 10 Hz | 0.01 | 0.011438 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 133787.0371 | 0.011368
10 mA, 20 Hz | 0.01 | 0.0099317 | 390 |  |  |  | 40 |  | 6980.6982 | 0.0099782
10 mA, 30 Hz | 0.01 | 0.0099905 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 295.2396 | 0.0099701
10 mA, 40 Hz | 0.01 | 0.0099955 | 160 |  |  |  | 40 |  | 220.2506 | 0.0099978
10 mA, 55 Hz | 0.01 | 0.0099992 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 11.5479 | 0.0099993
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 12.5827 | 0.01
10 mA, 1 KHz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 14.1418 | 0.01
10 mA, 5 KHz | 0.01 | 0.01 | 1000 |  |  |  | 70 |  | 8.165 | 0.01
10 mA, 10 KHz | 0.01 | 0.0099982 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 5.7746 | 0.0099982
100 mA, 10 Hz | 0.1 | 0.11546 | 700 |  |  |  | 40 |  | 123921.1136 | 0.12506
100 mA, 20 Hz | 0.1 | 0.099894 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 6707.1575 | 0.10009
100 mA, 30 Hz | 0.1 | 0.099917 | 390 |  |  |  | 40 |  | 73.7162 | 0.09969
100 mA, 40 Hz | 0.1 | 0.099947 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 356.3138 | 0.099907
100 mA, 55 Hz | 0.1 | 0.099983 | 160 |  |  |  | 40 |  | 18.933 | 0.099987
100 mA, 300 Hz | 0.1 | 0.099998 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 9.5745 | 0.099996
100 mA, 1 KHz | 0.1 | 0.1 | 160 |  |  |  | 40 |  | 23.6286 | 0.099997
100 mA, 5 KHz | 0.1 | 0.10001 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 4.9994 | 0.10001
100 mA, 10 KHz | 0.1 | 0.10002 | 0.25 |  |  |  | 200 |  | 14.1395 | 0.10002
1 A, 10 Hz | 1 | 1.0142 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 154550.9977 | 0.99954
1 A, 20 Hz | 1 | 0.99684 | -1 |  |  |  | 40 |  | 6158.5934 | 1.0045
1 A, 30 Hz | 1 | 0.99863 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 389.2384 | 0.99816
1 A, 40 Hz | 1 | 0.99918 | 640 |  |  |  | 40 |  | 353.6565 | 0.9989
1 A, 55 Hz | 1 | 0.99968 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 15.0048 | 0.99968
1 A, 300 Hz | 1 | 0.99982 | 640 |  |  |  | 40 |  | 9.576 | 0.99983
1 A, 1 KHz | 1 | 0.99977 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 15.0035 | 0.99979
1 A, 5 KHz | 1 | 0.99971 | 800 |  |  |  | 70 |  | 8.1673 | 0.9997
1 A, 10 KHz | 1 | 0.99911 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 5.0045 | 0.99911
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.01 | 300.97 | 126.42 | 210.02 | 20.5 | 20.61